## ملخص العطاء العام
: للمواصفات الفنية الخاصة بوحدة التبريد (ثلاجة تبريد و نظام طاقة شمسية على مستوى المديرية بتمويل من صندوق الأمم المتحدة للسكان لسنة 2026م ) UNFPA 2026 Technical Specifications for 1 Refrigerator and Solar System – Districts
ملخص العطاء العام
م No. | بيان إجماليات قيمة الأعمال | الكمية المطلوبة | المبلغ (رقما) دولار تنتقل تلقائيا من جدول الكمياتت | الإجمالي | المبلغ ( كتابتا ) $ Total Writing USD
Lot 1 | Medical Refrigerator ثلاجات | 55 | 0 | 0
Lot 2 | Solar System for 1 Refrigerator – 24 Hours | 55 | 0 | 0
الإجمالي Total |  |  |  | 0
نسبة التخفيض إن وجدت ( % ) Reduction Percentage if any |  |  |  |  | يتم ادخال نسبة التخفيض ان وجدت
مبلغ التخفيض بعد ضربة بالاجمالي العام |  |  |  | 0
الاجمالي الكلي بعد التخفيض Total Total after Reduction |  |  |  | 0
الإجمالي الكلي كتابتا مبلغ وقدره: يكتب بالاسفل على النقاط The grand total, written in amount and capacity .......................................................................................................................................................................................................
تعليمات هامه وارشادات وعلى المتقدم في المناقصة الإقرار بها
أقر أنا الموقع أدناه، مقدم العطاء، بأنني قمت قبل تقديم الأسعار بما يلي: 1) دراسة جميع وثائق العطاء، بما في ذلك شروط تقديم العطاء، صيغة العطاء، الشروط العامة والخاصة للتعاقد، الرسومات الهندسية، جداول الكميات، والمواصفات الفنية الخاصة بالمشروع. 2) دراسة احتياجات المشروع من المواد والمعدات والأعمال المطلوبة، وطرق الحصول عليها، والأسعار السائدة والمتغيرة خلال فترة التنفيذ. 3) تسعير كل Lot وفقاً لجداول الكميات والمواصفات الفنية المرفقة، ويُعد السعر المقدم شاملاً لكافة التكاليف اللازمة للتوريد والتنفيذ والتركيب والتشغيل والتدريب والتسليم، بما في ذلك النقل، العمالة، الأدوات، المعدات، المواد، الملحقات، الاختبارات، الضمان، وأي تكاليف أخرى لازمة لإنجاز وتسليم الأعمال بصورة مكتملة وجاهزة للاستخدام، ما لم يرد خلاف ذلك صراحة في وثائق العطاء. 4) أقر بأن منظمة ديم للتنمية تحتفظ بحق الترسية لكل Lot بشكل مستقل، وليس بالضرورة أن يتم التعاقد مع نفس المورد/المقاول لتنفيذ جميع الـ Lots. 5) أقر بأن أي خصم مقدم في هذا الملخص، إن وجد، سيتم احتسابه على إجمالي العرض أو على الـ Lot المحدد بحسب ما يوضحه مقدم العطاء في عرضه، ولن يتم قبول أي تعديل لاحق على الأسعار بعد فتح العطاءات إلا وفق الإجراءات المعتمدة.
مقدم العرض \ Supplier :
الاسم / Name : |  |  |  |  | التوقيع / Signature :
التاريخ / Date : |  |  |  |  | رقم التلفون / Phone Num:
الختم / Stamp :
مكتب الصادق للعمارة والهندسة

## جدول الكميات والمواصفات
المواصفات الفنية الخاصة بوحدة التبريد (ثلاجة تبريد و نظام طاقة شمسية على مستوى المديرية بتمويل من صندوق الأمم المتحدة للسكان لسنة 2026م ) UNFPA 2026 Technical Specifications for 1 Refrigerator and Solar System – Districts
ITEM NO | Specifications المواصفات | The price for one system (per cooling unit). | Specifications of the submitted offer. | Country of origin. | Brand Name | Warranty. | Other Comments
General Instructions to Suppliers
1 | Fridge 360L (Quantity 1)
Specifications
Capacity: 330–360 liters.
Temperature range: between +2°C and +8°C.
Temperature display resolution: 0.1°C; control accuracy around ±1°C or better.
Digital Temperature Control
Multi Air flow cooling technology
Interior fan for forced air circulation and uniform cooling.
Inverter compressor or energy efficient compressor technology.
Cooling Feature: Frost Free
Automatic defrost system
Refrigerant: R600a or approved equivalent
Lockable, Single glass door
Adjustable shelves
Number of shelves : minimum 3 shelves
Transparent interior / glass door suitable for visual stock inspection.
Door Gasket: Removable door gasket, antibacterial and antifungal where available.
Internal LED lamp.
High and low temperature audible and visual alarm
Power supply: 220V, 50Hz.
Rated power: approximately 189–230 W.
Estimated energy consumption: approximately 4.54 kWh/day or better.
Supplied with standard accessories, user manual, and warranty.
Application: medical, pharmaceutical, laboratory, and RH cold chain storage.
The refrigerator shall be brand new, suitable for continuous operation, and supported by local after sales service and spare parts availability.
2 | Solar System for Fridge 360L (1) 24 Hours
2.1 | Inverter (Quantity 1)
Main Technical Specifications
Waveform Type: Pure Sine Wave
Battery Voltage: 48V DC (48VDC)
Charge Controller: Built in 100A MPPT controller
Maximum PV Input Power: from 5000W to 6000W
MPPT Voltage Range: 85V ~ 450V DC
Battery less Operation: Supports operation with or without batteries (battery use is recommended)
IP from 24 to 66
Parallel Capability: Supports parallel connection of up to 6 units for higher capacity
Peak Efficiency: 93% – 94%
Communication Interfaces: USB / RS485 / CAN / WiFi supported
Key Features:
High Performance: Suitable for homes and small commercial installations.
Advanced Protection: Includes Type II SPD protection for the DC side.
LCD Display: Provides detailed information about the system’s status.
2.2 | Solar PV Modules:Quantity: 3 pieces.(Three panels per cooling unit)
Module capacity should not be less than 590 to 700 W at STC
Must conform to IEC 61215, 61730, 61701, 62716, ISO 9001, ISO 14001, and ISO 45001 certificates or equivalent.
The PV manufacturer should be approved as tier 1;
The PV modules junction box not less than IP68 & with 3 diodes.
Module Voltage: Not less than 1500 VDC (IEC).
VOC: 0.25% /C° or less.
Nominal operating cell temperature (NOCT): Not more than 45°C.
Tolerance of maximum power rating: 0 5 W.
Operating temperature: 40C° to 85C°.
Weatherproof DC rated MC4 connector & Fully Secured, not allowing for any loose connections & PID resistance.
Certificates and Data sheet of PV module that contains the P V & I V Curves, material and thickness of the frame and glass layer, all electrical and mechanical Data, Dimensions and Module area should be provided by bidders.
Warranty of high Performance
2.3 | Lithium ion Batteries (LiFePO4) - Quantity 1 piece.
Type: Low Voltage (LV) Lithium Iron Phosphate (LiFePO4) battery.
Nominal Energy: 5120 Who per piece.
Nominal Capacity: 100 Ah per piece.
Nominal Voltage: 51.2 V.
Operating Voltage Range: 40V - 58.4 V.
Max. Continuous Charge/Discharge Current: 100 Amps.
Min. Parallel Quantity: 2 batteries or more
Cycle Life: More than 5000 cycles at 25°C.
Installation Method: Can be wall mounted or floor mounted. Battery Chemistry:
LiFePO4.Supplying and installing solar panel holders to achieve the best solar.
2.4 | Galvanized, Solar Panel Mounting Structures Supply and installation of steel supports for mounting solar panels, designed to match the panel’s size and weight, coated with an anti-corrosion finish. The item includes welding, fixing, and installation works
2.5 | Double-Insulated DC Solar Cable (per meter): Unit price per meter Supply and installation of double-insulated DC solar cable per meter, resistant to UV radiation, humidity, and weather conditions, and compliant with TUV or IEC standards for solar applications. The cable shall be suitable for the system and shall include red and black cables as required, along with all terminals, connectors, and accessories necessary for installation.
2.6 | AC/DC Solar Protection Panel Supply and installation of a solar combiner panel with separate DC and AC sections for safe and organized interconnection of PV/inverter outputs.
2.8 | Installation Accessories (Lump Sum) Supply of all installation accessories required to complete the solar system, including PVC or metal trunking with full accessories, original MC4 solar connectors, cable ties/clamps, conduits, fasteners, bolts, brackets, nuts, and all materials required to complete the installation properly and safely.
Totals per Solar System for one Fridge Cooling System اجمالي سعر منظمومة الطاقة الشمسية الواحدة لوحدة تبريد واحدة |  | 0
مقدم العرض \ Supplier :
الاسم / Name : |  |  |  |  | التوقيع / Signature :
الختم / Stamp