## compiled list
# | Item | Item Description and Specifications | Unit of measurment | Quantity | Estimated Unit Cost (if known) | Estimated Total | Comment
Lab Equipment
1 | Laboratory Microscope, Binocular | Laboratory Microscope, Binocular With Binocular Observation Tube, Mechanical Stage, Condenser and Filter Holder. Objectives Magnification: 4X, 10X, 20X, 40X & 100X Achromatic Objective lines Eyepieces:Two eyepeices 10X widefield & 2 eyepiece 16X. Stand Focusing Mechanism: built-on graduated mechanical stage fine adjustments: single lateral with coarse and coaxial adjustment graduated both sides Square stage Type Illumination: Condenser type: focusable condenser Low voltage Light source: LED illumination System ( Easy to replace the lamp ) Tension Adjustment: Coaxial, X-Y controls Full range objectives magnification "Complete with: 2 pcs. Ground glass filter. 3 pcs. spare lamps 2 pcs, 150 ml Bottle of immersion oil 1 pc. Dust cover. 1 Cabinetwith/case 1 Cleaning brush" "Supplied With Complete Accessories. • Guarantee Spare Parts: (Ensure the supply of spare part for a Period of (7-10) years from the supply date of the device)" "Power Supply:- - 220- 240 V ~ ±10%, 50/60 Hz , 6V, 20W Bulb Illumination System or LED type" | Ea | 2 |  | 0
2 | Semi-Automated biochemistry Analyzer | Semi-Automated biochemistry Analyzer CONFIGURATION Benchtop Connection to LAS track Processing modes System capacity THROUGHPUT, max tests 50-70 /hr SAMPLE TYPE Serum/plasma Urine CSF SAMPLE SIZE, 2- 30 µL TEST MENU, min Acid phosphatase ,Alanine transaminase (ALT) ,Albumin ,Aldolase,Alkaline phosphatase (ALP) ,Amylase ,Aspartate transaminase (AST) ,Bilirubin (direct/total) ,Blood urea nitrogen (BUN) ,C3 ,C4 ,Calcium ,Chloride ,Cholesterol ,Cholinesterase ,Creatine kinase (CK) ,Creatinine ,CK-MB ,CO2 ,CSF protein ,GGT ,Glucose ,IgA ,IgG ,IgM ,Inorganic phosphorus ,Iron ,Lactate dehydrogenase (LDH) ,Magnesium ,P-5-P ,Potassium Total protein Sodium Thyroxine (T4) Tyroxine uptake (TU) Triglycerides Urea Uric acid Urine protein Others HBDH, lipase, neonatal bilirubin, RF, LD-1, ammonia, TIBC, LDL etc according to the device range of tests offered. PROGRAMMED 50 TESTS USER-DEFINABLE TESTS METHOD USED End point, rate, ISE OR Kinetic Optical system (Photometric ) Light source (Tungsten-halogen ) REAGENTS, TYPE Liquid Substitution Refrigerated onboard SYSTEM FEATURES Closed-tube sampling Direct sampling Liquid-level sensing Clot detection Auto dilution Abnormal values flag Auto verification Auto quality control Auto calibration DATA MANAGEMENT Display LCD Results stored Computer interface Current LIS vendor interfaces Printer Bar-code reader LINE POWER, 220 VAC, 50Hz Backup WATER REQUIREMENTS, <15 L/hr Onboard supply Supplied with: Set of 3 sterilisation containers adapted to the equipment proposed Water distiller device Operation Manual Service Manual Shall comply with CE marked or FDA Certificate. NOTE All equipment needing consumables must allow the possibility to use generic and/or locally made consumables and/or disposables. Compliance to this condition must be declared here by the bidders. Service/ maintenance and technical documents inclusive of schematics, component diagrams, trouble shooting , spare parts ordering information etc. 2 Years comprehensive warranty, from the date of installation and commissioning Installation & Commissioning must be done by manufacturer engineer Users training, by Specialist from the Suplier. The system offered shall be designed to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. | Ea | 3 |  | 0
3 | CBC Hematology Analyzer 3 Part | • Open system. • 21 parameters (WBC, WBC 3 diff , HTC, RBC, Hb, hematocrit, MCV , MCH, MCHc, RDW SD/RDW-CV, PLT, MPV, Pt Crit, PDW, PLCR ), with 3 Histogram WBC, RBC, PLT. • Impedance methods for WBC, RBC, PLT. Cyanide free reagent for HGB test . • low reagent consumption. • 2 reagent required for test 1 reagent for cleanser. • Sample volume required 18 µl for WB test and 20µ for PD test. • Input wide color touch LCD screen, key board and mouse . • Built-in thermal printer and support external printer. • Memory for 1000 results incl. histograms. • 60 sample per hour. • Auto sleep function. • two mode for test ( whole blood , pre-dilution ) • manual calibration and auto calibration ( calibrator, fresh blood ) • Built-in QC record program. • working temperature +5 C°~ +35 C° • Wide power supply AC 100-240, 50/60 HZ. • at least 4 USB port , 1 LAN port , HL7 protocol. • English language for system is essential. • Suitable online UPS. • Two years warranty. * Provision of reagents for three months • The installation and training for the staff should be performed by qualified BME engineer. • Reagents: All the reagents required for 1000 tests should be supplied with the equipment . • FDA,CE and any quality certificate. • Catalogue, Service and operation manuals (original and copy) to be provided. • The spare price list of all spares and accessories (including minor) required for maintenance and repairs in future after guarantee/warranty period should be attached. | Ea | 1 |  | 0
4 | Centrifuge 12 tube, complete w/accessories | Specification Speed regulation range for centrifuge tubes 100-3000 rpm (1610 × g) Digital time setting 1 - 90 min (increment 1 min) Timer sound signal + Rotor imbalance diagnostics (automatic stop, “IMBALANCE” warning) + Display LCD, 2 x 16 signs Nominal operating voltage 220 V, 50/60 Hz | Ea | 1 |  | 0
5 | Micropipette (100-1000 µl) | One of each size: volume micropipettes: 100-1000) - Variable Volume Micropipettes feature built-in tip ejectors and autoclavable tip cones. - Should work on a click-stop digital system, are easy to calibrate and maintain, and easy to disassemble for autoclaving. - Manufactured as per ISO 9001:2008,. Each pipette should be individually calibrated according to ISO 8655 standards - Calibration certificate must be provided with each pipette. - Accuracy and Precision values should be those laid down in the ISO 8655 standards. - Built-in, streamlined tip ejector facilitates easy tip ejection and access to narrow necked bottles and tubes. - Variable Volume Pipette: 100-1000 - Increments [µl]: 5 and the other is 1 - Volume [µl]: 100 to 1000 - Accuracy (±) %: 0.9 to 0.6 - Precision (±) %: 0.3 to 0.2 - It should be CE mark | Ea | 1 |  | 0
6 | Micropipette (10-100 µl) | One of each size: volume micropipettes: (10-100 µl). - Variable Volume Micropipettes feature built-in tip ejectors and autoclavable tip cones. - Should work on a click-stop digital system, are easy to calibrate and maintain, and easy to disassemble for autoclaving. - Manufactured as per ISO 9001:2008,. Each pipette should be individually calibrated according to ISO 8655 standards - Calibration certificate must be provided with each pipette. - Accuracy and Precision values should be those laid down in the ISO 8655 standards. - Built-in, streamlined tip ejector facilitates easy tip ejection and access to narrow necked bottles and tubes. - Variable Volume Pipette: 10- 100 µl - Increments [µl]: 5 and the other is 1 - Volume [µl]: 10 to 100 - Accuracy (±) %: 0.80/1 .003.00 - Precision (±) %: 0.15/0.40/1.50 - It should be CE mark | Ea | 1 |  | 0
7 | Lab Chairs | Backrest,Seat: PU leather or Micro Fiber Leather, High Density Rebound Foam, superior injection plastic board. Travel Distance: 45# Steel, Surface treatment: Chrome Plated or Electrostatic sparyed. Diameter of Five Star Base: High Polised Aluminum Alloy. (Max Capacity: 750KG) | Ea | 2 |  | 0
8 | Medical Lab DDA Cupboard | Product Features Steel construction Scratch & chip resistant epoxy powder coated Lighting on the inside and outside 1 cabinet with key lock alongside 2 adjustable glass shelves internally 2 fixed shelves and 1 upper compartment Technical Specifications Approx. Dim (mm): 715(W) x 180(D) x 760(H) Available in WM600-BZ c/w a battery operated buzzer | Ea | 2 |  | 0
Surgical sets/Equipment
9 | Caesarean Section Set | CAESAREAN SECTION SET Each Set Contents of: Tray Kidney shape with Lid S/S 200 X 30 mm 2 Tray Instrument with Lid 225X 125 X 50 mm Deep 2 Forceps Artery Spencer Wells Straight BJ.S/S 125mm 12 Forceps Artery Spencer Wells Straight BJ.S/S 200mm 2 Forceps Dissecting Plain serrated jaws S/S 150 mm 2 Forceps Dissecting Treves, 1x2 Teeth 150 mm 2 Forceps Tissue Allis fine 4 x 5 Teeth BJ.S/S 150 mm 2 Forceps Tissue Little wood 2 x 3 Teeth (bj).S/S 190 mm 3 Forceps Towel Holding cross Action S/S 100mm 6 Forceps Sponge Holding Rampley BJ S/S 240 mm 4 Handle size 3 for Scalple blades size 10-15 1 Handle size 4 for Scalple blades size 20-24 1 Needle Holder Mayo Hegar BJ.S/S 180mm 2 Scissors Mayo Straight SJ. S/S 200mm. 1 Scissors Mc Indoe/Metzenbaum CO Flat SJ.S/S 190mm 1 Retr. Doyen Ring Handle blade 95 mm Length 280 mm 1 Retr. Lange beck H/H 13x44 Blade, 210 mm Long 1 Retr. Morris Double-ended (52x38)&(52x64) mm 1 Pin Safety Mayo, forceps holding, 120mm 3 Suction Tube Pool's S/S 225mm 1 Forceps Haemostasis Green Armitage 200 mm 6 Sterilizing and storing case St. St. with filters 260x150x50 mm 2 | Set | 2 |  | 0
10 | Delivery Set | Kit contents: 1 x Scissors,Mayo,140mm,cvd,b/b 2 x Episiotomy, Barnes 126mm , cvd 1 x Scissors,gyneco,200mm,cvd,b/b 1x forceps, tissue, spring-type, 1 x 2 teeth, 14.5 cm 1x needle holder, Mayo-Hegar, straight, narrow jaw, 18 cm 2 x Forceps,artery,Kocher,140mm,str 1xBowl, ROUND, 100 ml, 80 x 35 mm, stainless steel 2x Kidney tray, 200x30mm 1 x Tray with cover large size (compatible size with the items) Vaginal Speculum (Small, medium and large) Stainless steel, high quality material | Set | 3 |  | 0
11 | Circumcision set | 1xSterilizing Stainless Steel Box 28x14x6 cm 1xMayo Scissors Curved 14 cm 1xOperating Scissors B/B Straight 14 cm 1xCircumcision Clamp Infant (8 mm) 1xFlat Handle Scalpel Concave blade 50 mm 1xFlat Handle Scalpel Convex blade 50 mm 4xBackhaus Towel Forceps 12 cm 2xMurphy‐Pean Forceps Straight 13 cm 1xDissecting Forceps No Teeth 14 cm | set | 2 |  | 0
12 | Forceps Hartmann Crocodile | Forceps Hartmann Crocodile 14cm (Reusable Autoclavable Stainless Steel) x 1 Crocodile Hartmann ENT Forceps - 5.5 inch | Ea | 2 |  | 0
13 | IUD insertion set | Gallabins Uterine Sound 12.5” Cusco Vaginal Speculum small Cusco Vaginal Speculum medium Cusco Vaginal Speculum large Rampley Sponge Holding Forceps 9.5” Luer Vulsellum Forceps 1:1 teeth, 9” Spencer Wells Artery Forceps 8” straight IUCD Removal Hook Laminated guide to caring for your instruments high quality stainless steel | set | 3 |  | 0
14 | Jar Dressing ( Big) | Round sterilized drums for autoclave and hot air sterilizer. • Material: Stainless Steel AISI304 Cylindrical containers used to sterilize dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave) or hot air sterilizer and to keep them as "sterile. Drums have an effective closing lid with a clip lock system, a carrying handle, air vents system to allow steam to circulate freely during the sterilization cycle/process | Ea | 2 |  | 0
15 | Jar Dressing (Medium ) | Round sterilized drums for autoclave and hot air sterilizer. • Material: Stainless Steel AISI304 Cylindrical containers used to sterilize dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave) or hot air sterilizer and to keep them as "sterile. Drums have an effective closing lid with a clip lock system, a carrying handle, air vents system to allow steam to circulate freely during the sterilization cycle/process | Ea | 2 |  | 0
16 | Instrument Tray with Lid including dressing kit | Dressing set, 8 pieces. Set contents: - Kidney Dish 20 cm (8") - medium - Gallipot 150 ml - 6 oz - medium - Kocher artery forceps - straight - 16 cm - Dressing Forceps (Thumb) - 16 cm - English TOE dissecting forceps - plain - 14 cm - Mayo scissors - straight - 15 cm - Rampley sponge holding forceps - str. - 18 cm - Instrument tray with lid 25 x 30 x 6.2 cm | Ea | 2 |  | 0
17 | Cheatle Forceps With Jar | Stainless Steel Diameter: 2 inch Height: 7 inch | Ea | 2 |  | 0
18 | Scissors, Deaver, 140 mm, straight, sharp | Scissors, Deaver, 140 mm, straight, sharp . Stainless steal, good quality | Ea | 3 |  | 0
19 | Speculum,vaginal | Size: Large Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 16 - 18% chromium; 10 - 14% nickel; 2 - 3% molybdenum. | Ea | 3 |  | 0
20 | Medical Instrument trolley | High quality trolley, instrument, with shelves, s/s, large, 900 (l) x 600 (w) x 800 (h) mm. Tubular or rectangular frame with two fixed stainless steel shelves, 150 mm swivel castors. Frame made of 25 mm round or rectangular tubing of stainless steel. Thickness of tubing 2 mm. With push handle. All stainless steel sheet metal of 1,5 mm thickness. Size: approx 900 (l) x 600 (w) x 800 (h) mm. | Ea | 2 |  | 0
Radiology
21 | CR system without printer | Complete CR system Computed Radiology must be a state of the art system manufactured by a reputed brand or manufacturer adhering to following specifications. CR system should broadly comprise of following modules/ components: a) Image recording system (cassettes & reading plates) b) Image reading system (reader/ digitizer) c) Identification & CR processing workstation. d) Dry imager / Printer. 1. Image recording system (cassettes & imaging plates). The following sizes of radiography cassettes along with image plates should be supported by the unit : a. 35 cm X 43 cm or 14” X 17” : 1nos. b. 24 cm X 30 cm or 10” X 12”: 1nos. c. 18 cm X 24 cm or 8” X 10”: 1 nos. 2. Image reader (CR reader/ digitizer) a) The CR reader / digitizer should be able to process 65 image plates/hr or more of the largest size cassette b) CR reader / digitizer must be able to handle phosphor image plates. CR reader capable of handling latest Dual side /needle/structured/columnar image plates will be preferred. c) It should have a resolution of 5 (minimum) to 10 (maximum) pixels/mm for standard resolution cassettes . d) Digitiser must have a resolution of 20 pixel / mm (minimum) . e) Gray scale resolution: CR reader / digitizer should have a minimum resolution of 12bits/ pixel for images sent to CR processing station. 3. Identification Station & processing server a) The processing station latest version have 2GB RAM DDR, 500 GB HDD (or more) in RAID configuration and 14 inch or more clinical grade monitor (Medical monitor). b) Processing server capable of identification of patient demographics to the acquired images will be preferred, else a separate identification station must be provided. c) The server and/or ID station must be DMWL (DICOM modality worklist) compliant to access patient and study data from hospital information system (HIS) or Radiology information system (RIS). d) It should provide display of acquired images with greater details of demographics viz. patient/ study listing for easy access e) The server must provide full amount of post processing features viz. geometric corrections, window level algorithms, annotation like markers, predefined text, drawing lines and geometrical shapes, multi-scale image processing, measuring distance and an f) It should facilitate full fledged DICOM printing and should be able to print multiple formats of patient study. g) Should be able to send DICOM images to DICOM workstation or PACS without loss of information h) Should be equipped with DICOM CD writer for transferring image i) Should be able to store image on external device viz. CD or pen drive etc. j) The system should have a facility to indicate over/under exposure in the preview screen. Kindly specify the image preview time. k) The software must have dedicated peadiatric image processing. 5.Supplied with : a) Suitable UPS back up must be provided for 30 minutes backup for the whole system c) Operation Manual d) Service Manual Service/ maintenance and technical documents inclusive of schematics, component diagrams, trouble shooting , support remote service (2 year free), spare parts ordering information etc. Installation & Commissioning two Years comprehensive warranty, from the date of installation and commissioning Users training, by Specialist from the Supplier. Shall comply with FDA Certificate or CE mark. | Ea | 3 |  | 0
Other
22 | Table,gynaeco,delivery,w/access | Technical specifications: Gynecological examination and delivery table, 3 sections, including accessories Mounted on 4 sturdy supports, all finished with height adjustable feet All sections fitted with non-removable padded upholstery Robust mechanics allow for manual repositioning of all sections between gynecological and obstetric use Transfer bars connect all lower distal portions of the 4 supports, providing maximal structural strength Back section: Adjustable via secured pawl and gear ratchet, safe for patient and operator Pelvic section: Can be tilted to Trendelenburg position with lever handle Sides of this section are fit with vertical handgrips, length 20 cm Padded knee crutches are height and width adjustable, set with robust clamps with heavy knob Fixing of the crutch holders is solid steel and welded to the frame of the bed This section integrates a support for a slide-out basin-tray Leg section: Recesses entirely downwards, 90 degrees When elevated and fully extended, all sections align to perfectly flat surface. Materials: High resistance to corrosion (tropical environment) Frame: epoxy coated tubular steel Adjustable feet: rubber or nylon Sliders/fixtures for the knee crutches: tubular steel, welded to the bed frame Padded upholstery: high-density polyurethane foam, density 27-33 kg/m3 Cover: plastic, flexible highly tear resistant, anti-static, flame retardant, disinfectant- and liquid proof, washable Basin-tray: Austenitic stainless steel, 18/10 Dimensions: All 3 sections extended, including padding: 162-198x72-88x72-88cm (l x w x h) Frame: 2.7-3.3cm (outside, across), 1.8-2.2mm (thickness) Padded upholstery: 4.5-7.5cm (h) Carrying capacity: 135-165 kg Slide-out basin-tray, capacity: 4L | Ea | 2 |  | 0
23 | Wheel Chair | Metal construction, Nicle chrome plated, foldable, with two wheels 25-inch dia,and two 8 inch dial, and hand brakes, aluminium foot rest, seat and back made from Lather, adult size. High resistance to corrosion (tropical environment). | Ea | 2 |  | 0
24 | Portable Compressor Nebulizer Machine For Adults &Pediatric with accessories | air flow at least 5 L/min. Minimum drug nebulisation flow greater than: 0.2 ml/min. Integrated drug reservoir with a minimum capacity: 6 ml. Residual volume (not operating below) less than 2 ml. Median Mass Aerodynamic Diameter (MMAD) of particles 2.44 µm. Noise level ≤ 55 dB Sturdy construction, suitable to be disinfected with hospital-grade products. Power requirements: 240 Volts - 50/60 Hz SUPPLIED WITH Instructions for assembly, use and maintenance in English. 1 x carry bag (depending on model supplied). 2 x reusable mouthpiece. 2 x air tubing set. 4 x reusable adult face masks. 4 x reusable paediatric face masks. 5 x air filters (depending on the model supplied). 2 x nebuliser tank. | Ea | 2 |  | 0
25 | Sphygmomanometer, Mucury, with stethoscop | Desk Type Mercury Sphygmomanometer Cuff Type Blood pressure Usage/Application Hospital Color Black Features : heavy duty and good quality Finishing Fine Features: High efficiency Dimensionally accurate | Ea | 10 |  | 0
26 | Examination bed with wheels | Examination table/bed, in 2 sections Mounted on 4 sturdy supports, all finished with height adjustable feet Both sections fitted with non-removable padded upholstery Backrest angle adjustable via secured pawl and gear ratchet, safe for patient and operator When fully extended, both sections align to perfectly flat surface Materials: High resistance to corrosion (tropical environment) Frame: epoxy coated tubular steel Adjustable feet: rubber or nylon Padded upholstery: high-density polyurethane foam, density 27-33 kg/m3 Cover: plastic, flexible highly tear resistant, anti-static, flame retardant, disinfectant- and liquid proof, washable Carrying capacity: minimum 160 kg Dimentionsvapproximatly: 190 cm Lx 65-70 cm W x 80-90 cm H. Metal | Ea | 1 |  | 0
27 | Ward Screen 3 Sections (Bedscreen,hospital,on castors) | Material: High quality metal construction, painted with electrostatic powder Multiple leg supports for stability. Extendable / Retractable. White epoxy coated frame or Stainless steel for durability and ease of cleaning. complete with pvc curtain Dimensions: 170 cm(h) x 190 cm(w) x 50 cm(d) Mobile on 50-75mm dia. swivel castors | Ea | 4 |  | 0
28 | Mechanical baby scales with sliding weights | Technical Data: Heavy duty, high quality Capacity:35 lbs, 16 kg Graduation (g):5 g, 1/4 oz Measure (W x H x D):550 x 180 x 290 mm , 21,7 x 7,1 x 11,4 inch Net weight:6,4 kg , 14,1 lbs Functions and Properties:Reset-to-zero function | Ea | 4 |  | 0
29 | Adult Weight Height Scale | TECHNICAL SPECIFICATIONS: Disply type: mechinal hand Measuring range: from 5 up to a minimum of 180kg. Minimum graduation: 100g Reading time in seconds: < 15 seconds Readout in metric (kg) Display: easily readable in low light working situations, black coloured numbers on silver surface. Can be read from both sides. With easy zero adjustment possibility. With stabilizing mechanism for faster reading of results Adjustable feet allow for horizontal levelling. Height measuring rod included | Ea | 4 |  | 0
30 | Oxygen Cylinder (40) lit. with Regulator & humidifier | Capacity (Litres) :40 Litre Working Pressure :150 kgf/cm2 Type of Cylinder :Filled Cylinder Material :Iron Body Cylinder Height :5 feet Cylinder Diameter :220 mm Flow range of the regulator: From 0.15 L/min to 15 L/min Maximum possible flow: 35 L/mn above range Input pressure: Not less than 3.5 bar +/- 0,7 Suitable for all types of fitting: wall outlet, trolley. Autoclavable, unbreakable Clear vision of bubbles Humidification at 4 L/mn - 20° C: ~3 ml/h Humidification at 12 L/mn - 20° C: ~9 ml/h Capacity: 100 ml (distilled water) Porosity of diffuser: No more than 70 microns Humidifier, Autoclavable, for multipurpose uses in Medical and Surgical Sections. Shall have the capability to be associated with Flowmeters (O2) or Air , O2 Reserve Cylinders and O2 Generators Fixation to flowmeters or any compatible devices shall be easy and fast (1/4 twist shall be acceptable) Supplied With 1Tubing 3 Face Mask diffrent Size | Ea | 1 |  | 0
31 | Examination bed | Examination table/bed, in 2 sections Mounted on 4 sturdy supports, all finished with height adjustable feet Both sections fitted with non-removable padded upholstery Backrest angle adjustable via secured pawl and gear ratchet, safe for patient and operator When fully extended, both sections align to perfectly flat surface Materials: High resistance to corrosion (tropical environment) Frame: epoxy coated tubular steel Adjustable feet: rubber or nylon Padded upholstery: high-density polyurethane foam, density 27-33 kg/m3 Cover: plastic, flexible highly tear resistant, anti-static, flame retardant, disinfectant- and liquid proof, washable Carrying capacity: minimum 160 kg Dimentionsvapproximatly: 190 cm Lx 65-70 cm W x 80-90 cm H. Metal | Ea | 2 |  | 0
32 | Doppler,FHR detector,w/access | Foetal Heart Rate (FHR) doppler detector, with speaker, handheld, battery powered, with accessories. Technical Specification: Capable of detecting foetal heart rates (FHRs) in the range of 50 - 210 bpm, with 1 bpm resolution and 2 bpm accuracy. Probe (transducer) with 2 MHz frequency with a probe attached via a cable. The cable can be extended to a 25 cm length when stretched. Probe detector head diameter at least 39 mm. Handheld, weights less than 0.45 kg with probe and batteries. Built-in speaker with volume adjustment. LCD screen displays FHR, pulse indicator and battery status Visual alerts for status, system error and low battery Operates on 3 x AA batteries. Battery life is suffiecient for 10 hours continuous use. | Ea | 2 |  | 0
33 | 40 Liter Automatic Autoclave Steam Sterilizer | Sterilizing chamber volume: 40L Maximum working pressure:0.22Mpa Maximum working temperature: 134 C Adjustment of temperature: 105-134C Timer: 0-99min Heat average: ≤±1C Power: 2.7kW Voltage: 220V/50Hz Package size(mm): 690*860*700 G.W: 113kg | Ea | 1 |  | 0
34 | ECG Machine | Digital 12 Channel ECG Technical Features: Display: 8" TFT, touch screen, led backlight Lead: Standard 12 leads Patient leak current: <10 micro A Input Impedance: >50ΜΩ Frequency response: 0.05Hz ~ 150Hz (-3dB) Dimension of recording paper : 210mm (W) x 20m (L) high-speed thermal paper Paper speed : Auto record : 25mm/s, 50mm/s, error : ±5% Measurement parameters : HR, P-R. interval, P Duration, QRS Duration, T Duration, Q-T interval, Q-Tc, P Axis, QRS Axis, T Axis, R(VS), S(VI), R(VS) + S(VI) Dimension : 334mm (L) x 320mm (W) x 85m (H) Net weight : 6.2 Kg Memory : 1 ECG cable, 1 ground wire, 1 set limb electrode clamp 1 paper roll, 1 set chest electrodes, Power connection USB connection, Software (same languages as internal software) Manual: GB, FR, IT CMRR : > 60dB, > 100dB (Adding Filter) Time constant : Time constant > 3.2s Input Way : Floating and defibrillation protection CMRR : > 60dB, > 100dB (Adding Filter) EMG interference filter : 35 Hz (-3dB) Recording way : Thermal printing system Rhythm record : 25mm/s, 50mm/s, error : ±5% Manual record : 5mm/s, 6.25mm/s, 10mm/s, 12.5mm/s, 50mm/s, error : ±5% Sensitivity selections : 5, 10, 20, 40mm/mV, error : ± 5%. Standard sensitivity is 10mm/mV ± 2mm/mV. Auto record : Record setup according to auto record format and mode, automatically changing leads, measuring and analyzing. Rhythm record : Record setup according to rhythm record format and mode, automatically measuring and analyzing. Manual record : Record setup according to manual record format, manually changing leads. Product safety type : Class I CF applied part; there is defibrillation and packing protection circuit in it. Enduring polarization voltage : ± 500mV Noise level : <15μVp-p Power supply / Frequency : AC : 100 – 240V, 50/60Hz DC:14.8V, 3700mAh rechargeable lithium battery | Ea | 1 |  | 0
35 | Donation Blood Chair | Blood donation chair. Mounted on 4 sturdy supports, all finished with height adjustable feet. Fixed height. All sections fitted with non-removable padded upholstery. Robust mechanics allow for manual repositioning of all sections. Self-locking adjustable armrest. Materials: High resistance to corrosion (tropical environment). Frame: epoxy coated steel Adjustable feet: rubber or nylon. Cover: plastic, flexible, highly tear resistant, anti-static, flame-retardant, non-absorbing, waterproof and cleanable with hospital-grade disinfection products. Padded upholstery: high-density polyurethane foam, density 30 kg/m3. | Ea | 1 |  | 0
[5 empty] | Total | 0

## DP
HF Name/ Items | Laboratory Microscope, Binocular | Semi-Automated biochemistry Analyzer | CBC Hematology Analyzer 3 Part | Centrifuge, complete w/accessories | Micropipette (100-1000 µl) | Micropipette (10-100 µl) | Lab Chairs | Medical Lab DDA Cupboard | Caesarean Section Set | Delivery Set | Circumcision set | Forceps Hartmann Crocodile | IUD insertion set | Jar Dressing ( Big) | Jar Dressing (Medium ) | Instrument Tray with Lid including dressing kit | Cheatle Forceps With Jar | Scissors, Deaver, 140 mm, straight, sharp | Speculum,vaginal | Medical Instrument trolley | Table,gynaeco,delivery,w/access | Wheel Chair | Portable Compressor Nebulizer Machine For Adults &Pediatric with accessories | Sphygmomanometer, Mucury, with stethoscope | CR system without printer | Examination bed with wheels | Ward Screen 3 Sections (Bedscreen,hospital,on castors) | Mechanical baby scales with sliding weights | Adult Weight Height Scale | Oxygen Cylinder (40) lit. with Regulator & humidifier | Examination bed | Doppler,FHR detector,w/access | 35 Liter Automatic Autoclave Steam Sterilizer | ECG MACHINE | Donation Blood Table
Al-Luhoom Health Complex | [6 empty] | 2 | [13 empty] | 1 | 1 | 1 |  |  |  | 4
Al Buraiqa Health Complex | 1 | 1 |  |  | 1 | 1 | [18 empty] | 1
Al Memdarah Health Center | [19 empty] | 2 |  | 1 |  | 3 |  | 1 |  | 2 | 2
Dar Sad HC | 1 | 1 | 1 |  |  |  |  | 2 | [16 empty] | 1
Al Basateen HC |  | 1 | [18 empty] | 1 |  | 1 | 2 | [5 empty] | 1 |  |  | 1
Asheikh Othman Health Complex | [10 empty] | 2 | 2 | 3 | 2 | 2 | 2 | 2 | 3 | 3 |  |  |  |  | 5 | 1 |  |  | 2 | 2 |  | 1 | 1
CEmONC Al Buraiqa |  |  |  | 1 |  |  |  |  | 2 | 3 | [20 empty] | 1 | 1 |  |  | 1
Ash Sha'ab HC | [33 empty] | 1
Total | 2 | 3 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 3 | 2 | 2 | 3 | 2 | 2 | 2 | 2 | 3 | 3 | 2 | 2 | 2 | 2 | 10 | 3 | 1 | 4 | 4 | 4 | 1 | 2 | 2 | 1 | 1 | 1