## Specifi
Date:27/8/2024 |  | Date:23/2/1446
Tables of Technical Specifications for The Purchase of Medical Devices and Equipment to Equip Health Facilities in in Al-Qafr District, Ibb Governorate, Supported by Tamdeen Youth Foundation
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
TMY-1 | Item Description | Ambu Bag Resuscitation Set (infant)
Qty |  | 6
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark | Required Certificate.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must Have Lattest Version
5.3 | Quality: | High Quality,
6 | Type: | Resuscitator, Hand-Operated, Infant (Body Weight <5 Kg), Set.
7 | Materials: | Latex-free
7.1 | Materials life: | All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions
7.2 |  | made of transparent silicone rubber
7.3 | Capacity: | 280 ml
8 | Compressible Self-Refilling Ventilation Bag: | Silicone rubber.
9 | Non-rebreathing Patient Valve with Pressure Limiting Valve: | Polycarbonate/polysulfone/silicone
10 | Inlet Valve with Nipple for O2 Tubing: | Polycarbonate/polysulfone
11 | Oxygen Reservoir Bag: | Plastic.
12 | Masks: | Transparent silicone rubber.
13 | Accessories:
13.1 |  | Compressible self-refilling ventilation bag, capacity: 220-280ml
13.2 |  | Oxygen reservoir bag complete, capacity: 600ml
13.3 |  | mask 0 , face Round type, size infant 30-50 mm small
13.4 |  | mask, 1 face Round type, size infant 50-65 mm small
13.5 |  | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly
13.6 |  | Carrying case
14 | Estimated life-Span of the Item: | At Least 10 Years
15 | Warranty: | 1 Year.
16 | Other Specification: | Please Specify Other Specifications.
TMY-2 | Item Description | Ambu Bag / Resuscitation Set (Child)
Qty |  | 6
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark | Required Certificate.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must Have Lattest Version
5.3 | Quality: | High Quality,
6 | Type: | Resuscitator, Hand-Operated, Infant (Body Weight More Than 5 Kg), Set.
7 | Materials life: | All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions
8 |  | Made of transparent silicone rubber
9 | Capacity: | 700 ml
10 | Compressible Self-Refilling Ventilation Bag: | Silicone Rubber.
11 | Non-rebreathing Patient Valve with Pressure Limiting Valve: | Polycarbonate/polysulfone/silicone
12 | Inlet Valve with Nipple for O2 Tubing: | Polycarbonate/polysulfone
13 | Oxygen Reservoir Bag: | Plastic.
14 | Masks: | transparent silicone rubber.
15 | Accessories:
15.1 |  | Compressible self-refilling ventilation bag, capacity: 500-700ml
15.2 |  | Non-rebreathing patient valve with or without pressure limiting valve, patient connector outside/inside diameter (OD/ID): 22/15mm
15.3 |  | Oxygen reservoir bag complete, capacity: 2000-2600ml
15.4 |  | mask, face shape type, size Child small
15.5 |  | mask, face shape type, size Child medium
15.6 |  | mask, face shape type, size Child large
15.7 |  | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly
15.8 |  | Carrying case
16 | Life Span for Item: | 7 Years
17 | Warranty: | 1 Year.
18 | Other Specification: | Please Specify Other Specification
TMY-3 | Item Description | Ambu Bag Resuscitation Set (Adult)
Qty |  | 6
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Type: | Resuscitator, hand-operated, adult (body weight over 30kg), set.
7 | Materials: | Latex-free
7.1 | Materials life: | All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions
7.2 |  | Made of transparent silicone rubber
7.3 | Capacity: | 2000-2600 ml
8 | Compressible Self-Refilling Ventilation Bag: | Silicone Rubber.
9 | Non-rebreathing Patient Valve with Pressure Limiting Valve: | Polycarbonate/Polysulfone/Silicone.
10 | Inlet Valve with Nipple for O2 Tubing: | Polycarbonate/polysulfone
11 | Oxygen Reservoir Bag: | Plastic.
12 | Masks: | Transparent silicone rubber.
13 | Accessories:
13.1 |  | Compressible self-refilling ventilation bag, capacity: 1475-2000ml
13.2 |  | Non-rebreathing patient valve with or without pressure limiting valve, patient connector outside/inside diameter (OD/ID): 22/15mm
13.3 |  | Oxygen reservoir bag complete, capacity: 2000-2600ml
13.4 |  | Mask, face shape type, size adult small/teenager
13.5 |  | Mask, face shape type, size adult medium
13.6 |  | Mask, face shape type, size adult large
13.7 |  | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly
13.8 |  | Carrying Case
14 | Life Span for Item: | 7 Years
15 | Warranty: | 1 Year.
16 | Other Specification: | Please Specify Other Specification
TMY-4 | Item Description | Midwife Kit
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | Designed: | Heavy Duty Designed.
5.2 | Quality: | High Quality
6 | Each Set Contents of :
6.1 |  | 1X Stethoscope binaural complete.
6.2 |  | 1 X Sphygmomanometer aneroid 300mm w/cuff
6.3 |  | 1 X Book exercise 200x150mm ruled 8mm lines
6.4 |  | 1 X Pen ballpoint blue
6.5 |  | 1 X Sheeting plastic clear pvc 91cmx180cm
6.6 |  | 1 X Bag polyethylene selfـsealing 125x200mm
6.7 |  | 3 X Gauzeـpad sterile 12 ply 76x76mm long
6.8 |  | 2 XCotton wool sbsorbent nonـsterile 100g
6.9 |  | 2 X Poucii clear polypropylene 250x380mm long
6.1 |  | 2 X Tape umbilical nonـsterile 3mm widex 100m
6.11 |  | 1 X Bottal n/m round screwـcap 50ml amber glass
6.12 |  | 1 X Bottal w/m round screwـcap 50ml amber glass
6.13 |  | 1 X Bottal dropping 10ml amber glass
6.14 |  | 1 X Tapes measure 1.5m vinylـcoated fiberglasses
6.15 |  | 1 X Themometer clinical subـnormal 30 to 40 C.
6.16 |  | 1 X Nail clipper/file
6.17 |  | 1 X Scale spring baby 5kg 25g graduations
6.18 |  | 10 X Gloves surgeon’s latex size 7.5
6.19 |  | 1 X Apron utility 900mmx1mm opaque plastic
6.2 |  | 1 X Catheter tracheal deled w/glass mucus trap
6.21 |  | 1 X Stethoscope fetal pinard monaural
6.22 |  | 1 X Case for midwife kit with lid empty aluminum
6.23 |  | 1 X Nethyllergometrine table 0.125 mg 30
6.24 |  | 2 X Scissors surgical straight 145 mm b/b S.
6.25 |  | 2 X Forceps Hemostat Straight Rochester 160mm
6.26 |  | 1 X Bowl sponge 600 ml stainless steel
6.27 |  | 1 X bowl kidney 825ml (28oz) stainless steel
6.28 |  | 1 X Towel huck 430x500
6.29 |  | 1 X Soap toilet 113g bar unwrapped
6.3 |  | 1 X Soap box2 piece hinged plastic
6.31 |  | 1 X Brush hand surgeon’s with nylon bristles
6.32 |  | 1 X Thermometer clinical oral/rectal1
6.33 |  | 1 X Foil weeps to prevent neonatal loss 122x92 mm
8 | Estimated life-Span of the Item: | At Least 10 Years
9 | Warranty: | Minimum of 1 Year.
10 | Other Specification: | Please Specify Other Specification
TMY-5 | Item Description | Oxygen Cylinder 40L with Oxygen Regulator and Humidifier
Qty |  | 5
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality,
6 | Type: | Oxygen Cylinder Empty.40L
7 | Medical Oxygen Cylinder:
7.1 |  | Oxygen cylinder Spindle Key for Bull Nose, Spare Spindle Key for Pin Index
7.2 | Pressure: | Working Pressure 150Bar/200 Bar
7.3 | Capasity: | 40 Liters (Bull Nose British Standard).
7.4 | Material: | Special Steel (Anti Rust). Blue Or Black Color.
8 | Oxygen Trolley on castor:
8.1 | Quality : | High quality.
8.2 | Material: | Special Steel.
8.3 | Load: | (45-50 kgs cylinder).
9 | Oxygen Regulator and Humidifier:
9.1 | Flow Range: | From 0.15 L/min to 15 L/min.
9.2 | Maximum possible flow: | 35 L/mn above range.
9.3 | Input pressure: | Not less than 3.5 bar +/- 0,7
9.4 | Suitable for all Types of Fitting: | Wall outlet, trolley.
9.5 |  | Autoclavable, Unbreakable
9.6 |  | Clear Vision of Bubbles
9.7 | Humidification at 4 L/mn - 20° C: | ~3 ml/h
9.8 | Humidification at 12 L/mn - 20° C: | ~9 ml/h
9.9 | Capacity: | 100 ml (distilled water)
9.1 | Porosity of diffuser: | No more than 70 microns
9.11 |  | Humidifier, Autoclavable, for multipurpose uses in Medical and Surgical Sections. Shall have the capability to be associated with Flowmeters (O2) or Air,O2 Reserve Cylinders and O2 Generators
9.12 |  | Fixation to flowmeters or any compatible devices shall be easy and fast (1/4 twist shall be acceptable)
10 | Oxygen Masks (for Child.&Adult) &Connections :
10.1 | Quality : | High Quality.
10.2 | Material: | Silicone and Rubber.
11 | Accessories:
11.1 |  | 3 X Face Mask diffrent Size for adult. Resuable Type
11.2 |  | 3 X Face Mask diffrent Size for children. Type
11.3 |  | 1 X Tubing.
11.4 |  | 1 X Oxygen Trolley for Oxygen Cylinder 10 L
12 | Estimated life-Span of the Item: | At Least 10 Years
13 | Warranty: | Minimum of 1 Year.
14 | Other Specification: | Please Specify Other Specification
TMY-6 | Item Description | Oxygen Mask for Adult
Qty |  | 12
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Should be of Latest Model and Must Have Lattest Technology
5.3 | Quality: | High Quality,Safety, Ease of Use
6 | Genaral Specification:
6.1 |  | Mouth Piece Fits Perfectly With Ultra Soft Edges
6.2 |  | Designed For Comfortable Fitting With Soft Material
6.3 |  | Elastic Strap Ensures A Better Fit
6.4 |  | Light Weight Design
6.5 | Size: | Adult Type
6.6 | Material: | Non toxic
6.7 |  | Sterile, Single Patient Use.
7 | Accessoris:
7.1 |  | Supplied With Long Tubing
8 | Estimated life-Span of the Item: | At Least 3 Years
9 | Warranty: | Minimum of 1 Year.
10 | Other Specification: | Please Specify Other Specifications.
TMY-7 | Item Description | Oxygen Mask for Pediatric
Qty |  | 12
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Should be of Latest Model and Must Have Lattest Technology
5.3 | Quality: | High Quality,Safety, Ease of Use
6 | Genaral Specification:
6.1 |  | Mouth Piece Fits Perfectly With Ultra Soft Edges
6.2 |  | Designed For Comfortable Fitting With Soft Material
6.3 |  | Elastic Strap Ensures A Better Fit
6.4 |  | Light Weight Design
6.5 | Size: | Pediatric Type
6.6 | Material: | Non toxic
6.7 |  | Sterile, Single Patient Use.
7 | Accessoris:
7.1 |  | Supplied With Long Tubing
8 | Estimated life-Span of the Item: | At Least 3 Years
9 | Warranty: | Minimum of 1 Year.
10 | Other Specification: | Please Specify Other Specifications.
TMY-8 | Item Description | Ward Screen 4 Sections
Qty |  | 6
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must Have Lattest Version
5.3 | Quality: | High Quality,
6 | Design & Quality: | Bed screen, hospital, on castors.
7 | Construction: | Metal construction,painted with elecrostatic powder,4 screens / 4 folds complete with Pvc, mobile castors.
8 | General Description: | Bed screen, hospital, on castors.
9 | Technical Specifications:
9.1 |  | Hospital bed screen in 4 sections with curtains
9.2 |  | Bed screen mounted on swivel castors; positioned in broad stance for good stability.
10 | Frame: | Epoxy or enamel coated or chrome plated tubular steel.
11 | Curtains: | Plastic non-transparent, washable and flame retardant.
12 |  | Extendable / Retractable
13 |  | Multiple leg supports for stability.
14 |  | Hygienic, easy to clean panels.
15 | Castors: | Mobile on 50-75mm dia. swivel castors
16 |  | complete with pvc curtain.
17 | Guiding Dimentions: | 180-200 cm(h) x 200-220 cm(w) x 45-50 cm(d)
18 | Estimated life-Span of the Item: | At Least 7 Years
19 | Warranty: | Minimum of 2 Years.
20 | Other Specification: | Please Specify Other Specification
TMY-9 | Item Description | Suture Instruments and Dressing Set
Qty |  | 6
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material: | 304 Stainless Seel Or Better
7 | Each Set Contents of:
7.1 |  | 1 X Tray Instrument with Lid 225X 125 X 50 mm Deep
7.2 |  | 6 X Forceps artery Kocher COF BJ.S/S 180mm.
7.3 |  | 1 X Forceps Dissecting Bonney's 1X 2 Teeth S/S 150 mm
7.4 |  | 1 X Forceps Dissecting Plain serrated jaws S/S 150 mm
7.5 |  | 2 X Forceps Dressing Bryants serrated jaws S/S 125mm
7.6 |  | 4 X Forceps Towel Holding cross Action S/S 100mm
7.7 |  | 2 X Forceps Sponge Holding Rampley BJ S/S 240 mm
7.8 |  | 1 X Handle size 3 for Scalple blades size 10-15
7.9 |  | 1 X Needle Holder Kilner's 165 mm
7.1 |  | 1 X Scissors Mayo Straight SJ. S/S 200mm.
7.11 |  | 1 X Scissors Iris straight, Sharp points 115 mm
7.12 |  | 1 X Sterilizing and storing case St. St. with filters 230x130x50 mm
8 | Life Span for Item: | 10 Years
9 | Warranty: | Minimum of 1 Year.
10 | Other Specification: | Please Specify Other Specifications.
TMY-10 | Item Description | Aneroid Sphygmomanometer
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 |  | A device designed to measure blood pressure consisting of an inflatable cuff that fits around a limb (arm or thigh),
7 |  | An inflation bulb for controlling the air pressure within the cuff, An aneroid manometer, and tubing.
8 |  | The aneroid manometer consists of a metal bellows, which expands as the pressure in the cuff increases, and a mechanical amplifier that transmits this expansion through a lever to an indicator needle, which rotates around a circular, calibrated scale.
9 | Type: | The manometer is hand held (Portable).
10 |  | Composed of a cuff containing an inflatable bag.
11 |  | The inflatable bag is connected via a tube to a bulb with an integrated manometer needle gauge.
12 | Cuff Material : | The cuff is made of durable material (e.g. nylon), which is non-deformable, and washable at 30°C.
13 |  | The cuff is fitted with double Velcro fastening, enabling a tight and secure fit around arms.
14 |  | The cuff is reinforced at both sides.
15 |  | The bag is inflated by means of the flexible bulb connected via a tube.
16 | Material Tube: | Silicone Rubber, Crack Resistant.
17 | Length tube between: | 50 to 70 cm, Or Better.
18 | Gauge Graduated: | 0 - 300mmHg (min) in 2 (max) mmHg increments, with pressure release valve.
19 | Accuracy : | (+/- 3mm Hg.)
20 | Accessories:
20.1 |  | 1 X Size cuff for adult .
20.2 |  | 1 X Size Cuff for Child.
20.3 |  | 1 X Size Cuff for Neonate.
20.4 |  | 1 X Plastic Protective Case
21 | Life Span for Item: | 10 Years
22 | Warranty: | Minimum of 1 Year.
23 | Other Specification: | Please Specify Other Specifications.
TMY-11 | Item Description | Stethoscope Binaural Bowels Stainless Steel
Qty |  | 6
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality,
6 | Type: | Stethoscope Binaural Bowels
7 | Quality: | Heavy Duty and High Quality
8 | Product Description:
8.1 |  | A Mechanical listening device designed for listening to sounds from the heart and lungs. Comprises a membrane atthe listening head connected by a split "Y" tube to the headgear with ear pieces that are placed into the users ears.
9 | Specifications:
9.1 |  | Double cup, dual-use (adult and pediatric auscultation) chest piece in stainless steel or chrome plated brass.
9.2 |  | Adult diaphragm 43-47mm; pediatric diaphragm 28-36mm.
9.3 |  | Y tube treated rubber or PVC with 8-11mm diameter.
9.4 |  | Sensitivity from 3.2dB to 26dB in a range from 50 to 1000Hz for cardiology.
9.5 |  | Sensitivity 8.1dB in a range from 600 Hz to 1,500Hz for pneumology.
9.6 | Arms: | Stainless Steel
9.7 |  | Removable Plastic Ear-Pieces.
9.8 |  | Latex-free
10 | Accessories:
10.1 |  | 4 X Ear caps made of elastomer rubber
10.2 |  | 1 X Storge Case
10.3 |  | 1 X Operation Manual
11 | Estimated life-Span of the Item: | At Least 10 Years
12 | Warranty: | Minimum of 1 Year.
13 | Other Specification: | Please Specify Other Specifications.
TMY-12 | Item Description | Digital Glucometer
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
5.2 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Digital Glucometer.
7 | Technical Specifications:
7.1 |  | Sample volume one drop
7.2 |  | Test Time Less than 10 seconds
7.3 |  | With Memory
7.4 |  | Set complete with prick pen
7.5 |  | Lancets
7.6 | Measurement Unit : | mmol/L Or mg/dL
7.7 | Sample Volume : | 0.5 microliters
7.8 | Display : | LCD Display
7.9 |  | Battery Operated
8 | Accessories:
8.1 |  | 500 X Glucose Strips
8.2 |  | 1 X Carry Bag.
8.3 |  | 2 X Battery (Spare part)
9 | Life Span for Machine: | 10 Years
10 | Warranty: | Minimum of 1 Year.
11 | Other Specification: | Please Specify Other Specifications.
TMY-13 | Item Description | Cheatle Sterilizer Forceps
Qty |  | 7
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | General Description:
6.1 |  | Cheatle Sterilizer Forceps
6.2 | Length: | 20-27 cm.
6.3 | Material: | 100% Stainless steel.
6.4 | Type: | Reusable.
7 | Life Span for Item: | 10 Years
8 | Warranty: | Minimum of 1 Year.
9 | Other Specification: | Please Specify Other Specifications.
TMY-14 | Item Description | Sterilizer Lifter
Qty |  | 5
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | General Description:
6.1 |  | Sterilizer Lifter
6.2 | Dimensions: | 20 x 15 x 3.5-5 cm
6.3 | Material: | Plastic
6.4 | Type: | Reusable.
7 | Life Span for Item: | 5 Years
8 | Warranty: | Minimum of 1 Year.
9 | Other Specification: | Please Specify Other Specifications.
TMY-15 | Item Description | Autoclave Tape , 2-5 cm X 50 m , Roll
Qty |  | 5
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | General Description:
6.1 |  | Autoclave Tape
6.2 | Length: | 50 m.
6.3 | Width: | 2-5 cm
7 | Life Span for Item: | 2 -3 Years
8 | Warranty: | Minimum of 1 Year.
9 | Other Specification: | Please Specify Other Specifications.
TMY-16 | Item Description | Sterilization Sheet , Roll Type
Qty |  | 7
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | General Description:
6.1 |  | Sterilization Sheet , Roll Type
6.2 | Length: | 100 m.
6.3 | Dimensions: | 60 x 60 cm
7 | Life Span for Item: | 2 -3 Years
8 | Warranty: | Minimum of 1 Year.
9 | Other Specification: | Please Specify Other Specifications.
TMY-17 | Item Description | Laboratory Microscope Binocular
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Machine should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Type: | Binocular
7 | Eyepieces: | Plese Specify
8 | Interpupillary distance adjustment, mm: | Adjustable
9 | Nosepiece Configuration: | Quadruple or better
10 | Objectives magnification: | 4x,10x; 40x; 100x (oil)
11 | Objectives Type: | Achromatic OR Plan achromatic
12 | Filter:
12.1 | Color(s): | Please Specify
13 | Illumination:
13.1 | Condenser Type: | Built-in
13.2 | Numerical Aperture: | Please Specify
13.3 | Light Source: | Halogen or Led
14 | Stand:
14.1 | Focusing Mechanism: | Moving Stage
14.2 | Coarse, Fine Adjustments: | Coaxial Coarse and Fine Adjustment
15 | Stage:
15.1 | Type: | Mechanical
15.2 | Tension adjustment: | Coaxial
15.3 | Motion: | X-Y
16 | Accessories:
16.1 | Adapter(s): | Required
16.2 | Dust cover: | Required
16.3 |  | 50 X Microscope Slides 76x26mm.
16.4 |  | 100 X Cover slides 22x22mm.
17 | Certification from the Manufacturer:
17.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
17.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
17.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
17.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
17.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
17.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
17.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
17.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
17.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
18 | Life Span for Machine: | 10 Years
19 | Maintenance:
19.1 |  | 2 Years Free Maintenace.
19.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
19.3 |  | Application Software and Interface Connection Included.
20 | Power Supply: | 100 to 240 V ~ ±10%, 50 Hz
21 | Other Specification: | Please Specify Other Specifications.
TMY-18 | Item Description | Tape Measure, Durable Fiberglass
Qty |  | 5
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Designed: | Heavy duty Designed
4.1 | Quality: | High Quality.
5 | Type: | Long-Lasting, Durable Fiberglass
6 | Size: | 150 Centimeters, Graduated in cm and inches
7 |  | Heat Sealed Graduations in Black and Red
8 |  | Steel Clip End Hooks
9 | Life Span for Machine: | 10 Years
10 | Warranty: | Minimum of 1 Year.
11 | Other Specification: | Please Specify Other Specifications.
TMY-19 | Item Description | Thermometer Digital
Qty |  | 24
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Power Type: | DC Complete.
7 | Quality: | Heavy duty and high quality
8 |  | For oral, underarm, or rectal use
9 | Operation Details: | Press ON/OFF button to activate. The unit would beep and display.
10 | Display: | This LCD display test will go on for around 2 seconds
11 | Low Temperature Display: | Temperature < 32°C Display L°C.
12 | High Temperature Display: | Temperature > 43°C Display H°C.
13 | Accuracy: | 35°C ----- 39°C ±0.1°C.
14 | Display Resolution: | 0.1°C.
15 | Normal Beep Sound: | 10 long beep sounds in 10 seconds.
16 | Fever Beep Sound: | (>37.8°C) : 30 short beep sounds in 10 seconds.
17 | Auto Power Off: | 8 minute ±2 minute
18 | Battery: | AG3 or LR41
19 | Life Of Battery: | 2 years or 2500 times.at least
20 | Accessories:
20.1 |  | 1 X Plastic Case.
20.2 |  | 2 X Spare Part Battery.
20.3 |  | Supplied With Complet Accessories.
21 | Life Span for Item: | 7 Years
22 | Warranty: | Minimum of 1 Year.
23 | Other Specification: | Please Specify Other Specifications.
TMY-20 | Item Description | Torch Pen Light For Examination
Qty |  | 5
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Power Type: | DC Complete.
7 |  | Torch Pen Light For Examination, Spot Light
8 |  | High-performance pen-lights for medical diagnosis.
9 |  | Available with LED illumination.
10 |  | Bright and natural white light by high-performance LED
11 | LED Life | At least 20.000 hours
12 | Material: | ABS casing with metal clip, disinfectable
13 |  | Easy to use with on/off switch
14 | Accessories:
14.1 |  | 1 X Case
14.2 |  | 3 X Spare Battery
14.3 |  | 2 X Spare Led.
15 | Life Span for Item: | 7 Years
16 | Warranty | Minimum of 1 Year.
17 | Other specification | Please Specify Other Specifications.
TMY-21 | Item Description | Dressing Trolley
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material of Construction: | Heavy duty Stainless Steel 304 or better and high quality
7 | General Description: | Dressing trolley with two non-removable shelves mounted on four swivel castors.
8 | Technical Specifications: | Guard rails on three sides on top shelf and Second Shelf
8.1 |  | 2 level shelves
8.2 | Carrying Capcity Approx: | 80-100 Kg
8.3 | Mobile On Swivel Castors : | 100mm (2 with brakes).
9 | Accessories:
9.1 |  | 1 x Dressing Drum Holder.
9.2 |  | 1 x Thermometer Holder.
9.3 |  | 1 x Forceps Container Holder.
9.4 |  | 1 x S.S. Dressing tray attached at the top. With Collapsible Holder
9.5 |  | 1 x Waste bin
9.6 |  | 1 x S.S. drawer below the first shelf.
9.7 |  | 1 x S.S Sponge Bowl.
9.8 |  | 1 x Dressing Jar.
10 | Guiding Dimentions of trolley : | 650-800Lx400-500 Wx900-1000mmH at least
11 | Life Span for Item: | 7 Years
12 | Warranty: | Minimum of 2 Years
13 | Other Specification: | Please Specify Other Specification
TMY-22 | Item Description | Vaccine Carrier
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan,USA and UK
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Vaccine Storage Capacity: | 3-5 litres
7 | Weight Fully Loaded: | 6-8 kg
8 | External Materials: | Woden Polyester Or Equivalent
9 | Internal lining materials: | Aluminum bag Or Equivalent
10 | Number of Icepacks Required: | (4-8)
11 | Life Span for Item: | 10 Years
12 | Warranty: | Minimum of 1 Year
13 | Other Specification: | Please Specify Other Specification.
TMY-23 | Item Description | Gynecology Examination Coach
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must have lattest Version
5.3 | Quality: | High Quality.
5 | General Description: | Examination Gynecology Coach, with Accessories
6 | Technical Specifications:
6.1 |  | Examination Gynecology Coach , 3 Sections
6.2 |  | Mounted on 4 sturdy supports, all finished with height adjustable feet
6.3 |  | All sections fitted with non-removable padded upholstery
6.4 |  | Robust mechanics allow for manual repositioning of all sections between gynecological and obstetric use
6.5 |  | Transfer bars connect all lower distal portions of the 4 supports, providing maximal structural strength
7 | Back Section:
7.1 |  | Adjustable via secured pawl and gear ratchet, safe for patient and operator
8 | Pelvic Section:
8.1 |  | Can be tilted to Trendelenburg position with lever handle
8.2 |  | Sides of this section are fit with vertical handgrips, length 20 cm
8.3 |  | Padded knee crutches are height and width adjustable, set with robust clamps with heavy knob
8.4 |  | Fixing of the crutch holders is solid steel and welded to the frame of the bed
8.5 |  | This section integrates a support for a slide-out basin-tray
9 | Leg Section:
9.1 |  | Recesses entirely downwards, 90 degrees
9.2 |  | When elevated and fully extended, all sections align to perfectly flat surface.
10 | Materials:
10.1 |  | High resistance to corrosion (tropical environment)
11 | Frame: | Stainlees steel
12 | Adjustable feet: | Rubber or nylon
13 | Sliders/fixtures for the knee crutches: | Tubular steel, welded to the bed frame
14 | Padded upholstery: | High-density polyurethane foam, density 27-33 kg/m3
15 | Cover: | Plastic, flexible highly tear resistant, anti-static, flame retardant, disinfectant- and liquid proof, washable
16 | Basin-tray: | Austenitic stainless steel, 18/10
17 | Dimensions: | All 3 sections extended, including padding: 180-200x85-100x70-90 cm (L xW x H)
18 | Carrying capacity: | 150-180 kg
19 | Slide-out basin-tray, capacity: | 4L
20 | Accessories:
20.1 |  | 1 X Complete set of tools required for assembly
20.2 |  | 2 x Leg holders with canvas straps, adjustable height and width
20.3 |  | 2 x Knee crutches, adjustable height and width
20.4 |  | 1 x Basin-tray
20.5 |  | 1 x Double Foot Steps.
21 | Estimated life-Span of the Item: | At Least 7 Years
22 | Warranty: | Minimum of 2 Years.
23 | Other Specification: | Please Specify Other Specification.
TMY-24 | Item Description | Bed Sheet
Qty |  | 100
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
4 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
5 | Design & Quality: | Heavy Duty, High Quality
6 | Material: | Sheet draw mixed fiber or Cotton.
7 | Dimensions: | 120*200 cm Or Better
8 | Estimated life-Span of the Item: | At Least 7 Years
9 | Warranty: | Minimum of 1 Year
10 | Other Specification: | Please Specify Other Specification.
TMY-25 | Item Description | Blanket , Adult
Qty |  | 10
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Design & Quality: | Heavy Duty, High quality
5 | Material: | Blanket Mixed fiber,Washable,Adult Plain Color with cotton case
6 | Dimensions: | 120-150*200-220 cm Or Better
7 | Estimated life-Span of the Item: | At Least 7 Years
8 | Warranty: | Minimum of 1 Year
9 | Other Specification: | Please Specify Other Specification.
TMY-26 | Item Description | Filing Cabinet Four Drawers
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Type: | Filing Cabinet Four Drawers
7 | Construction: | Made from entirely metallic , from matter Heavy-gauged , Epoxy-Painted and oven-baked components
8 |  | Recess Handle and Label Holder
9 |  | Accommodates All Types filing
10 |  | Each Drawer has High Sides for Additional Strength
11 |  | Locking with 2 Keys
12 | Height : | 110-120 cm at least
13 | Width : | 45-55 cm at least
14 | Depth : | 60-65 cm at least
15 | Life Span for Item: | 7 Years
16 | Warranty: | Minimum of 1 Year
17 | Other Specification: | Please Specify Other Specifications.
TMY-27 | Item Description | Cupboard 2-Doors Glass Storage
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must have lattest Version
5.3 | Quality: | High Quality,
6 | Material of Construction: | Heavy Duty Stainless Steel 304 or better.
7 | General Description:
7.1 |  | Cupboard Double Glass Door
8 | Technical Specifications:
8.1 |  | Recessed safety glass in the doors allows for viewing of cabinet content
8.2 |  | Safety glass is supported by door frame on both sides, leaving no glass edges exposed
8.3 |  | Inside reinforced fixtures facilitate height adjustment of the 5 Metal Removable Shelves
8.4 |  | Doors are triple hinged, closed with handle and lockable with key
8.5 |  | Two-point lock, slides in frame at the top and base
8.6 |  | Key is unique for each single cabinet, key-bow folds away from key-blade avoiding breakage when inside lock
8.7 |  | Clearance underneath the cabinet allows for easy cleaning
9 | Materials:
9.1 |  | High Resistance to Corrosion (Tropical Environment)
9.2 | Frame, side panels, base, top : | Metal
9.3 | Shelves: | Epoxy Coated Plate Steel
9.4 | Doors: | Hardened Safety Glass, Grounded Edges, With Key-Lock.
10 | Dimensions:
10.1 | Cabinet, overall: | 80-100 x 40-60 x 180-200cm (l x w x h).
10.2 | Clearance underneath: | 10 cm.
10.3 | Safety glass: | 3-5 mm (thickness).
10.4 | Front glass area: | 65-70%.
10.5 | Frame, drawer, door and panels: | 1.2-1.8 mm (thickness).
10.6 | Door, opening angle: | 180 degree
11 | Carrying capacity, each shelf: | 25-30 kg.
12 |  | Shall be dust & rust proofed
13 | Items Supplied With:
13.1 |  | 1 x Complete set of tools required for assembly
13.2 |  | 1 x Set of 5 shelves
13.3 |  | 1 x Set of 2 keys, unique per cabinet
14 | Estimated life-Span of the Item: | At Least 7 Years
15 | Warranty: | Minimum of 1 Year.
16 | Other Specification: | Please Specify Other Specifications.
TMY-28 | Item Description | Stainless Steel 3 Seats Hospital Waiting Chair
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material of Construction: | Heavy duty Shall have stinless steel and backrest .The structure shall be from wrought Iron or better and high quality
7 |  | Bench, Patients' waiting.
8 |  | Composition from one piece ( One bench).
9 |  | The backrest height to be fit with seat capacity.
10 | Indicative dimensions: | Length 2000mm. Backrest: 900mm ± 10%.
11 | Life Span for Item: | 7 Years
12 | Warranty: | Minimum of 1 Year.
13 | Other Specification: | Please Specify Other Specifications.
TMY-29 | Item Description | Single Chair With Armrest
Qty |  | 8
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Construction: | Frame Made from (Strong Plastic)
6.1 |  | frame made from (chorome – plated tubular steel with floor studs) , the seat and backrest made from strong plastic (stackble) ,the base made from chorome – plated tubular steel.
6.2 |  | The seat and backrest thick 0.5 cm at least
7 |  | Chair for waiting rooms and other purposes
8 |  | Sngle Shair
9 |  | The bench designer form healthful and convenient
10 | Height : | 65-75 cm at least
11 | Width : | 55-60 cm at least
12 | Depth : | 35-45 cm at least
13 | Estimated life-Span of the Item: | At Least 7 Years
14 | Warranty: | Minimum of 1 Year.
15 | Other Specification: | Please Specify Other Specifications.
TMY-30 | Item Description | Modern Wooden Office Table with Drawers
Qty |  | 5
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material: | Wooden.
7 | Type: | Office Desk small with double side drawer
8 | Composition: | The Office Desk consist of one table to form rectangle.
9 |  | contain 2 Pedestal with 6 drawers(3 drawers for each side) (Locks /Keys)
10 | Thick desk: | 2 cm at least
11 | Pedestal side :
11.1 |  | Count 3 drawers each side
12 |  | Ball bearing silent sliders. (Locks /Keys)
13 | Width : | 180 cm at least.
14 | Depth : | 71 cm at least.
15 | Height : | 72 cm at least.
16 | Life Span for Item: | 7 Years
17 | Warranty: | Minimum of 1 Year.
18 | Other Specification: | Please Specify Other Specifications.
TMY-31 | Item Description | Office Chair with Armrest Based on Five Castors.
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
6 |  | Modern Office Chair
7 | Material: | Back and seat shall be made of upholstered thick 5 cm at least , Rest arms made of upholstered
8 | Type: | Office chair with armrest based on five castors.
9 |  | Chair for office purpose , Single ,round (moving)
10 | the base: | Metal can moving in all directions.
11 | Castors: | Shall be mobile on 5 castors on 5 stars metallic stand.
12 | chair : | Can Moving up & down be Hydraulic system
12.1 |  | Tilting action thru decentered axis enabling feet resting on ground
13 | Height : | 100 cm at least
14 | Width : | 50 cm at least
15 | Depth : | 51 cm at least
16 | Life Span for Item: | 7 Years
17 | Warranty: | Minimum of 1 Year.
18 | Other Specification: | Please Specify Other Specifications.
TMY-32 | Item Description | Examination Couch
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 |  | Table, Examination Heavy duty, high quality
7 | Construction: | Epoxy coated tubular steel. Or Stainless Steel
8 | General Description: | Examination table
9 | Technical specifications:
9.1 |  | Examination table, in 2 sections
9.2 |  | Mounted on 4 sturdy supports, all finished with height adjustable feet
9.3 |  | Both sections fitted with non-removable padded upholstery
9.4 |  | Backrest angle adjustable via secured pawl and gear ratchet, safe for patient and operator
9.5 |  | When fully extended, both sections align to perfectly flat surface
9.6 |  | Transfer bars connect all lower distal portions of the 4 supports, providing maximal structural strength
10 | Materials: | High resistance to corrosion (tropical environment).
11 | Frame: | Epoxy coated tubular steel.
12 | Adjustable feet: | Rubber or Nylon.
13 | Padded upholstery: | High-density polyurethane foam, density 27-33 kg/m3.
14 | Cover: | Plastic, flexible highly tear resistant, anti-static, flame retardant, disinfectant- and liquid proof, washable.
15 | Dimensions:
15.1 |  | Examination tabel, two sections extended, including upholstery: 185-190 x 55-60 x 70-80cm (l x w x h).
15.2 | Frame: | 2.7-3.5cm (outside, across), minimum 1.5 mm thickness.
15.3 | Upholstery: | 4.5-5.8cm (h).
16 | Carrying Capacity: | 200-250 kg Or Better For Examination Couch
17 | Knockdown construction: | Included.
18 |  | 2 Foot Steps
19 | Design & Quality: | Sturdy 2 step level foot Steps for, examination Heavy duty, high quality
20 | Construction: | Epoxy coated tubular steel. Or Stainless Steel
21 | General Description: | Steps for Examination couch
22 | Technical specifications:
22.1 |  | Sturdy 2 step level
22.2 |  | Mounted on robust supporting legs spaciously arranged for maximal stability
22.3 |  | Steps with embossed anti-slip pattern, and feet finished with heavy duty anti-slip rubber ends
22.4 |  | Transfer bars connect all lower distal portions of the supports, providing maximal structural strength
23 | Materials:
23.1 | Anti-slip feet: | High resistance to corrosion (tropical environment)
23.2 |  | Rubber or Plastic
23.3 | Frame: | Epoxy Coated or Chrome Plated Tubular Steel
23.4 | Steps: | Heavy Duty Solid Plastic
24 | Dimensions:
24.1 | Footstool, overall: | 50-65x30-40x30-35cm (L x W x H)
24.2 | Frame: | 2.5-3.5 cm (outside, across), 1.5-1.8 mm thickness
25 | Carrying capacity: | 150-200kg Or Better For Foot Steps.
26 | Knockdown construction: | Included.
27 | Life Span for Item: | 7 Years
28 | Warranty: | Minimum of 2 Years.
29 | Other Specification: | Please Specify Other Specification.
TMY-33 | Item Description | Pilow / Cushions
Qty |  | 21
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Design & Quality: | Heavy Duty, High Quality
5 | Material: | Pillow Foam with Plastic cover
6 | Dimensions: | 40-50*20-30 cm Or Better
7 | Estimated life-Span of the Item: | At Least 7 Years
8 | Warranty: | Minimum of 1 Year.
9 | Other Specification: | Please Specify Other Specification.
TMY-34 | Item Description | ECG Machine 12 Ch. With Trolley
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality,Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | System: | Fully Automatic, Microprocessor Controlled, Channel Resting Electrocardiograph, Mains and Battery Operated, Built-in LCD
7 | Modes: | Suitable for Adult/Pediatric/Neonate modes, and easy to all functions and parameter settings. Alarm mute On/Off recorder, Built-in alarm mute, on/off freeze, on/off recorder, Leads Switching, automatic signal Calibration, and Leads-off indicator, Automatic Switch off: After 5 min inactive
8 | Sensitivity: | 2.5, 5, 10, 20 mm/mV Or Better.
9 | Frequency: | Range for diagnostic: 0.67-150 Hz, range for filtering: Any line frequency, low and muscle artifact/high frequency.
10 | Recording: | Channels: Multi channels, ECG leads: 12 standard leads, method: Automatic/manual, speed: 35, 65 mm/sec, automatic printout on standard paper printer if printer is connected, printout on internal thermal printer for realtime manual mode.
11 | Measurements: | PR, PQ, QT, ATC, P, QRS, T, HR; others by requirements and clinician preference. PR, PQ, QT, ATC, P, QRS, T, HR; others by requirements and clinician preference, Arrhythmia analysis.
12 | Visual Indicators: | Visual Indicators: Loose contact, lead fail, low battery, system status, artifacts and anthers.
13 | Patient leakage current: | Less than 10µ A.
14 | Battery | Included. At Least 4 Hrs.
15 | Certification from the Manufacturer:
15.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
15.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
15.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
15.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
15.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
15.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
15.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
15.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
15.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
16 | Life Span for Machine: | 10 Years
17 | Maintenance:
17.1 |  | 2 Years Free Maintenace.
17.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
17.3 |  | Application Software and Interface Connection Included.
18 | Accessories: | 3 x Roll Paper
18.1 |  | 1 X High Quality Trolley With Arm.
18.2 |  | 1 X Case, carrying.
18.3 |  | 2 X Packs, thermosensitive. 50mmx20m.
18.4 |  | 2 X Electrodes, Limb. clamp type.
18.5 |  | 2 Kits X Electrodes, Suction.
18.6 |  | 200 Pcs X Electrodes, Resting. Disposable.
18.7 |  | 10 X Adaptors, Clip.Alligator.4mm.
18.8 |  | 2 X 250g Gel, electroconductive.
18.9 |  | 2 X Patient cable. Banana plug.
19 | Power supply | 100 to 240 V ~ ±10%, 50/60 Hz.
20 | Other Specification: | Please Specify Other Specification.
TMY-35 | Item Description | Portable Doppler Ultrasound System , Diagnostic Machine
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 | Genaral Specification: | The Unit shall be safe to use both for -The Operator and the Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Type: | Portable System With Trolley on four castors , two with brakes.
7 | System technology: | Digital beam former , upgradeable Hardware & Software.
8 | Applications: | Abdominal , OBS/GYN,Urology. Cardiology ,Vascular , Small Parts.and More Applications Please Specify.
9 | Scanning Techniques: | 1) Electronic convex scanning. 2) Electronic linear scanning.
10 | Monitor: | High resolution B & W monitor with tilt and swivel and a size of 12- 15 inches or Better.
11 | Key Board: | Full key board with built in trackball.
12 | Display Modes: | B, B/B , M , B/M , 2D Imaging , Color Doppler mode , Harmonic Imaging , Pulsed Wave Doppler , 3D/4D.
12.1 | Other Display Modes: | Please Specify,
13 | Gray Scale Max: | 256 Levels Or Better.
14 | Dynamic Range: | 100 dB Or Better.
15 | Frame Rate, Frame /Sec: | Please Specify
16 | Image Control: | Gain , STC , Image processing , Image shift , Zoom and post processing.
17 | Processing Channels: | 256 Channels Or Better.
18 | M-Mode Sweep Speed: | 4 Steps Or Better.
19 | Presets: | 5 Or Better.
20 | Cine Memory: | Please Specify No. of Images.
21 | Image Storage: | On HDD & CD and USB , Please specify storage capacity.
22 | Measurement: | Distance , Area , Volume and Fetal Parameters.
23 | Calculation: | Calculation Packages for Abdominal , OB/GYN.
24 | Connector Probe: | Minimum Two.
25 | Probe Holder: | Included.
26 | Probes:
26.1 |  | Convex probe 2-5 - 5 MHz multi frequency (Priced probe)
26.2 |  | Vaginal probe 6.5 MHz multi frequency (Priced probe)
27 | Printer: | B & W Printer
28 | UPS: | Included Online Type (2 KVA)
29 | Accessories:
29.1 |  | 5 X Gel, 500 g
29.2 |  | 6 x Roll Paper
29.3 |  | 1 X Original Trolley
29.4 |  | 1 X Dust Cover
29.5 |  | 1 X Operation Manual
29.6 |  | 1 X Service Manual
30 | Power supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
30.1 | Rechargeable Battery: | Rechargeable battery with built in charger with indicators for charging and low battery
31 | Essential Requirement:
31.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
31.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
31.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
31.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
31.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
31.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
31.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
31.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
31.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
31.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
31.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
32 | Estimated life-Span of the Device: | At Least 10 Years
33 | Maintenance:
33.1 |  | 2 Years Free Maintenace.
34 | Training: | User /Nurses training, by Specialist from the Supplier.
35 | Other Specification: | Please Specify Other Specification.
TMY-36 | Item Description | Elisa Machine Reader & Washer
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 | Genaral Specification: | The Unit shall be Safe to Use
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | System: | Open system
7 | Elisa Reader:
7.1 |  | Two basic readers for 96-well micro plates, Flexibility and reliability, Fast reading: 5s for 96well plates, Internal memory holds up to 64 assay, Data manipulation through software, Spectral range: 340-750 nm, Read-out range: 0-3.5 Abs,
7.2 |  | Measurement range: 0-2.5 at 405 nm, Reading speed: 5 s entire plate, Shaking: Linear shaking , 3 speed . Photometric Methode,
7.3 |  | Optical System: measurement and reference Channels, multiple channels required, Measurement modes: Single/dual wavelength, endpoint, Singe wavelength kinetic and two point.
7.4 |  | Calculation modes: Ranged Grayscale Threshold, Floating cut-off, Point-to-point, Linear regression and Cubic spline Reading speed: 5 entire plate, linearity: {( -2 )–(+2) }% Abs. CV<0.2%.
7.5 |  | Resolution, OD: 0,01, Compatible Microplates: roun/flat bottom preferred, programmable Assays is preferred.
7.6 |  | with Display, interface.
7.7 |  | Supply with: all required accessories to put Machine in function, Cover, Adapter, fuse, bulbs and calibration Plate.
8 | Elisa Washer:
8.1 |  | Automated, Compatible plates: 96-well plates,
8.2 |  | Reliable and easy to use, Co-axial wash heads for trouble-free washing and low residual volumes 8 and 12 way interchangeable wash heads, In-line rinse capability automatic pump shutdown. Wash/waste bottle capacity: 2 L+ 2 L. Additional wash bottle capacity: 2 L. Priming consumption: 15-20 ml. Wash heads: 8-and 12 ways. Program cards: 1-4 washes.
8.3 |  | Accuracy: ˃2%, Max dispense volume (ml/well): ˂1
9 | UPS: | Required ( Online Type)
10 | Certification from the Manufacturer:
10.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
10.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
10.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
10.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
10.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
10.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
10.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
10.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
10.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
11 | Life Span for Machine: | 10 Years
12 | Maintenance:
12.1 |  | 2 Years Free Maintenace.
12.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
12.3 |  | Application Software and Interface Connection Included.
12.4 |  | Spare Parts list with code NO
13 | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
14 | Other Specification: | Please Specify Other Specification.
TMY-37 | Item Description | Blood Bank Refrigerator
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | General Description: | Blood Bank Refrigerator
7 | Maintains Internal Air Temperature: | Between +2°C and +6°C Or Better.
8 |  | Blood Bank Refrigerator Is used to store blood bags under controlled temperature.
9 |  | Temp Range-should have adjustable termperature control range from +2 °C to +6 °C, factory preset at 4 °C Or Better.
10 | Capacity: | Should accommodate 100-200 or More units blood and storage internal volume should be at least 200-300 liters.
11 | Refrigerator system:
11.1 |  | The System should have high density CFC -free urethane foam insulation to protect cabinet from ambient temperature fluctuation.
11.2 |  | The system should have positive, forced, air circulation to maintain temperature uniformity at all shelf levels, with quick recovery +/- 1 degree C
11.3 |  | The system should have sensors for activating automatic defrost cycle to minimize the frost build up.
11.4 |  | The system should have automatic condensate removal with no requirement for separate drainage lines.
11.5 |  | Must have fully extendable drawer slides.
12 | Construction: | Internal construction should be made up of high grade stainless steel (min 22 G). External construction Corrosion resistant sheet, at least I mm thickness.
13 | Internal Temp Control :
13.1 |  | System should have temperature control range from +2 °C to +6 °C.
13.2 |  | Temperature control resolution should be better than 0.1 degree C.
13.3 |  | Cooling down time of max of 120-150 min on half load.
13.4 |  | Should be able to have connectivity to computer and data logger.
13.5 |  | Door System should be lockable with double glass doors for better safety
14 | Saftey System:
14.1 |  | System should have large and clear Digital displays for the set/run parameters.
14.2 |  | The system should have weekly chart recorder for temperature changes
14.3 |  | The system should have key operated set point for the added security.
14 | Alarm:
14.1 |  | System should have audible/visual warnings for over-temperature, under temperature and power failure with visual status reports on critical functions.
14.2 |  | System should have battery backup and connections for remote alarm contacts.
14.3 |  | Should have adjustments for uneven bases. The adjustments should be easy to use like rotating a screw at the legs In the base.
14.4 |  | Scratch resistant internal lining of the cabinet (stainless steel or aluminium).
15 | Display: | Externail tempreture display should be attached, position adjustable
16 | Certification from the Manufacturer:
16.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
16.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
16.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
16.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
16.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
16.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
16.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
16.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
16.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
17 | Life Span for Machine: | 10 Years
18 | Maintenance:
18.1 |  | 2 Years Free Maintenace.
18.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
18.3 |  | Application Software and Interface Connection Included.
18.4 |  | Spare Parts list with code NO
19 | Accessories:
19.1 |  | 4-5 Removable Shelves.
20 | Line Power , VAC: | 100-240, 50/60 Hz
21 | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
22 | Other Specification: | Please Specify Other Specification.
TMY-38 | Item Description | Ventilator Machine
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Ventilator System:
6.1 |  | Microprocessor Controlled Used for Pediatric,Adult and Neonate.
6.2 |  | Active Humidifier.
6.3 |  | Compressor or turbine >60000 hour - full 7 years warranty
6.4 |  | Nebulizer
7 | Ventilation Mode: | Volume and Pressure Type of Breaths , Patient types Adult & Pediatric.and Neonate
7.1 |  | BIPASIC , VG and Bi level ventilation modes ,PC-CMV, PC-BIPAP1 / SIMV, PC-SIMV, PC-AC, PC-APRV, PC-PSV,VC-MMV, A/C, CPAP/PEEP, VC-AC,VC-CMV, VC-SIMV, SPN-CPAP/PS, SPN-CPAP/VS, SPN-CPAP, SPN-PPS)and newer modes for Adult & pediatric.and Neonate
7.2 | Invasive Ventilation Modes: | Please Specify.
7.3 | Non-Invasive Ventilation Modes: | Required With all Modes.
8 | Controls /Setting Ranges:
8.1 | Tidal Volume, ml: | 2-2000 ml Or Better
8.2 | Respiratory Rate, Breaths/Min: | 0-150
8.3 | Trigger Mechanism: | Pressure, flow
8.4 | FiO2, %: | 21-100
8.5 | Inspiratory Flow Rate, L/min: | 0-180
8.6 | Flow Pattern/Waveform Adjustment: | Required
8.7 | Inspiratory Pressure, cm H2O: | 0-80
8.8 | IE Ratio: | Please specify.
8.9 | Sigh Breath Function: | Preferred
8.1 | PEEP/CPAP, cm H2O: | 0-20
8.11 | Pressure Support, cm H2O: | 0-40
8.12 | Leak Compensation: | Required
8.13 | Auto 100%/Increase O2 Button: | Required
8.14 | Control Panel Lock: | Required
8.15 | Automatic Ventilation Mode: | Required
9 | Special Functions:
9.1 | Pediatric Application: | Required
9.2 | Neontel Application: | Required
9.3 | Nebulizer: | Required
9.4 | Apnea Pack Up Ventilation: | Required
9.5 | Pause (Plateau ): | Required
10 | Graphic Waveform , Trends & Loops:
10.1 |  | Flow
10.2 |  | Volume
10.3 |  | Pressure
11 | Ventilation Display / Monitored: | Multicolor LCD / LED display
11.1 |  | Peak inspiratory pressure
11.2 |  | Mean airway pressure
11.3 |  | PEEP pressure
11.4 |  | Tidal volume
11.5 |  | Minute volume
11.6 |  | Spontaneous Minute Volume
11.7 |  | FiO2 (Analyzed %)
11.8 |  | Respiratory rate
11.9 |  | Inspiratory time
11.1 |  | Expiratory time
11.11 |  | IE Ratio
12 | Alarm Limits (Adjustable):
12.1 |  | Low/high pressure limit
12.2 |  | Low O2
12.3 |  | Low / high PEEP / CPAP pressure
12.4 |  | Low / high tidal volume
12.5 |  | Low / high minute volume
12.6 |  | Low / high respiratory rate
12.7 |  | O2 Concentration : 18-100 %
12.8 |  | Apnea
12.9 |  | All alarms audible and visible
13 | Equipment Alarms:
13.1 |  | Power Failure
13.2 |  | Low Battery
13.3 |  | Disconnection
13.4 |  | Gas Supply Failure
13.5 |  | Flow Sensor Alarms
13.6 |  | Vent Inoperative
13.7 |  | Self-Diagnostics
13.8 |  | All Alarms: Audible and Visible
14 | Patient Alarms:
14.1 |  | Low/high FiO2
14.2 |  | Low minute volume
14.3 |  | High minute volume
14.4 |  | Low inspiratory pressure
14.5 |  | High pressure
14.6 |  | Loss of PEEP
14.7 |  | Apnea
14.8 |  | Continuous high pressure/occlusion
14.9 |  | Inverse IE Ratio
14.1 |  | High Respiratory Rate
14.11 |  | High PEEP
14.12 |  | Breathing circuit disconnect
15 | Reporting (vent alarms and patient status):
15.1 | View reports on display | Required
15.2 | Save data to USB | Required
15.3 | Send data via network | Required
16 | Inlet gas supply pressure (Air/O2): | (30-80) PSI
17 | On-Board Air Compresspr Or Turbine | Please specify
17.1 | Compressor for ventilator: | Medical type, high quality and heavy duty
17.2 |  | Original and compatible with the ventilator
17.3 |  | Water and oil free
17.4 |  | Low noise ( > 60 dB)
17.5 |  | Automatic compressor backup system .(Air compressor will switch on automatically when inlet pressure is failed )
17.6 |  | Easy accessible
17.7 | Turbine: | High Quality Turbine >60000 hour - Full 7 years Warranty Or More.
18 | Display : | Led, Touchscreen
19 | Display size, (in) : | Min 15 in
20 | Certification from the Manufacturer:
20.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
20.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
20.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
20.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
20.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
20.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
20.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
20.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
20.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
21 | Life Span for Machine: | 10 Years
22 | Maintenance:
22.1 |  | 2 Years Free Maintenance.
22.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
22.3 |  | Spare Parts List with Code No
23 | Accessories:
23.1 |  | 1 X O2 sensor (Permanent Type)
23.2 |  | 1 X Active Humidifier
23.3 |  | 2 X Reusable patient circuit for pediatric
23.4 |  | 2 X Reusable patient circuit for adult
23.5 |  | 2 X Reusable patient circuit for Neonate
23.6 |  | 1 X Resuable Test lung for Adult , Pediatric and Neonate
23.7 |  | 1 X Resuable Test lung for Pediatric and Neonate
23.8 |  | 3 X Disposable patient circuit for pediatric
23.9 |  | 3 X Disposable patient circuit for adult
23.1 |  | 3 X Disposable patient circuit for Neonate
23.11 |  | Nebulizer Build-In
23.12 |  | All Needed Accessories For Full Pediatric & Adult and Neonatal Use
24 | Training | User /Nurses Training, by Specialist from the Supplier. And Biomedical Engineers
25 | Internal Back-up Battery: | Operating Time, Hr : ≥2
26 | Self Diagnostic Error / Calibration Massage | Included
27 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single Phase
28 | Other Specification: | Please Specify Other Specification.
TMY-39 | Item Description | Patient Monitor With Holder
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Display size and Type: | 12-15 Inches Or Better. Touch Screen
High resolution multi-color display
Medical type (preferable) or supported with isolation transformer
7 | Displayed information: | Min. 6 vital waveforms, cab be colored separately
Numeric data for the measured vital parameters
Vital parameters (72) hrs trends
Vital parameters alarms audio visual
8 | Defibrillation Protection: | Available for the measured vital parameters
9 | Required vital parameters:
9.1 | ECG:
9.1-1 | Leads: | 3-5 Leads facility
9.1-2 | Values: | ECG lead waveform, label, HR gain as min.
9.1-3 | HR Range: | 30-200 bpm Or Better.
9.1-4 | Alarms: | Leads off, Hi + Low HR
9.1-5 | Gain: | 5, 10, 20 mm/mV Or Better.
9.1-6 | Preferable items: | ST, Arrhythmia, cascade ECG
9.2 | Respiration:
9.2-1 | Technology: | Impedance
9.2-2 | Values: | Respiration waveform, R.R.
9.2-3 | R.R. Range: | 5-80 bpm Or Better.
9.2-4 | Alarms: | Hi & low R.R.
9.2-5 | Apnea Alarm: | 15-25 sec. Preferable
9.3 | NIBP:
9.3-1 | Technology: | Oscillometric
9.3-2 | Values: | SYS , DIA , MEAN , P.R.
9.3-3 | Modes: | Auto, Manual
9.3-4 | Cuff Pressure Range:
9.3-4-1 |  | Up to 250mmHg Adult
9.3-4-2 |  | Up to 200mmHg Pediatric
9.3-4-3 |  | Up to 150mmHg Neonate
9.3-4-4 | Alarms: | Hi and low, SYS, DIA
9.4 | SPO2:
9.4-1 | Values: | SPO2 waveform, SPO2% , P.R.
9.4-2 | SPO2 range | 50 - 100 % Or Better.
9.4-3 | Pulse Rate Range: | 25-200 ppm Or Better.
9.4-4 | Alarms: | Hi & low SPO2 + P.R. sensor off
9.5 | Temperature:
9.5-1 | Values: | T1 and or T2
9.5-2 | Range: | 30-45°C Or Better.
9.5-3 | Alarm: | Hi & low
9.6 | ETCO2:
9.6-1 |  | Should be Main Stream capnography with display of CO2 and digital Values of EtCO2, FiCO2 & RR.
9.6-2 | EtCO2 Range : | 0-99 mmHg.
9.6-3 | FiCO2 : | 0 to 20 mmHg.
9.6-4 | Flow rate : | 50ml/min.
10 | Certification from the Manufacturer:
10.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
10.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
10.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
10.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
10.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
10.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
10.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
10.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
10.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
11 | Life Span for Machine: | 10 Years
12 | Maintenance:
12.1 |  | 2 Years Free Maintenace.
12.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
12.3 |  | Application Software and Interface Connection Included.
12.4 |  | Spare Parts list with code NO
13 | Required Accessories:
13.1 |  | 1 X 3/5 Leads wire ECG cable Adult.
13.2 |  | 1 X 3/5 Leads wire ECG cable Pediatric.
13.3 |  | 1 X NIBP reusable cuff large Adult
13.4 |  | 1 X NIBP reusable cuff pediatric
13.5 |  | 1 X NIBP reusable cuff Infant
13.6 |  | 1 X NIBP connection hoses
13.7 |  | 1 X SPO2 reusable finger probe + Extension cable adult.
13.8 |  | 1 X Temperature reusable sensor skin Type for Pediatric and Infant
13.9 |  | 1 X EtCO2 module with all accessories. ( Cable-Sensor)
13.1 |  | 1 X Original Holder In The Wall
13.11 |  | Operation Manual
13.12 |  | Service Manual
13.13 |  | Spare parts list with code NO.
14 | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
15 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
15.1 | Internal back-up battery | Operating time, hr : ≥2
16 | Other Specification: | Please Specify Other Specification.
TMY-40 | Item Description | Fetoscope - Audio Fetal Doppler
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | Genaral Specification: | The Unit shall be safe to use
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality,
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Type: | handheld.able for high load & hard work
7 | General Description:
7.1 |  | Foetal Heart Rate (FHR) doppler detector, with speaker, handheld, battery powered, with accessories
8 | Sensitivity From : | 10 - 12 weeks Or Better
9 | Technical Specifications:
9.1 |  | Foetal Heart Rate (FHR) ultrasound detector, single transducer
9.2 |  | Handheld lightweight detector
9.3 |  | Ultrasonic continuous doppler operating mode
9.4 |  | External transducer, connected to the main unit through a coiled wire cable
9.5 |  | Integrated transducer holed on the body of the main unit
10 | Ultrasound frequency : | 2 MHz ± 10 %
11 | Sensitivity : | 90 dB
12 | Ultrasonic Power Output : | mW, depending on the available model
13 | FHR Range : | 50 – 240 bpm ± 2 bpm, resolution 1 bpm
14 |  | Integrated loudspeaker and volume level adjustment
15 |  | Noise suppression feature
16 |  | Integrated headset connector
17 | Display: | LCD display, backlighted, with FHR, pulse, frequency, volume, battery and system status
18 | Alarms: | Audio/visual alarms for high/low FHR, battery status and system errors
19 |  | Automatic Self-Test
20 |  | Automatic Power-Off
21 | Accessories:
21.1 |  | 2 X 350 ml or 3 x 250 ml (depending on availability) bottles of US gel
21.2 |  | 2 X sets of commercial disposable alkaline batteries (separately packed)
21.3 |  | 1 X soft carry bag
21.4 |  | 1 X Instructions for assembly, use and maintenance in English
21 | Battery Autonomy: | >5 h or 600 exams, whichever higher
22 | Power Supply: : | 1.5 V AA (LR6) Commercial disposable batteries or 1.2 V Ni-MH commercial rechargeable batteries (not included), 2 or 3 units depending on the available model
23 | Accessories:
23.1 |  | 2 x 350 ml or 3 x 250 ml (depending on availability) bottles of US gel
23.2 |  | 2 x sets of commercial disposable alkaline batteries (separately packed)
23.3 |  | 1 x soft carry bag
23.4 |  | Instructions for assembly, use and maintenance in English
24 | Essential Requirement:
24.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
24.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
24.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
24.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
24.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
24.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
24.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
24.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
24.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
24.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
24.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
25 | Estimated life-Span of the Device: | At Least 5 Years
26 | Maintenance: | 2 Years free maintenance.
27 | Training: | User /Nurses training, by Specialist from the Supplier.
28 | Other Specification: | Please Specify Other Specification.
TMY-41 | Item Description | Electrical Nebulizer Ultrasonic Mobile
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | Genaral Specification: | The Unit shall be safe to use
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality,
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Nebulizing rate, mL/min:
6.1 |  | 3-5 mI / min At Least.
6.2 |  | Medical fluid chamber capacity = 250cc,medication cup 50cc .
6.3 |  | Nebulizing timer: 0 - 30 Minutes & Continuous.
6.4 |  | Crystal Operating Frequency : 1.7 Mhz
7 | Alarms :
7.1 |  | Door open.
7.2 |  | Occlusion.
7.3 |  | Low voltage.
8 | Type: | Mobile
9 | Operating pressure | 0-1.0 bar
10 | All standard accessories | All accessories (filters, tubings, face mask (Adult,Pediatric,Infant, etc.)
11 | Noise level: | Please specify
12 | Max. Pressure: | Please specify
13 | Certification from the Manufacturer:
13.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
13.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
13.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
13.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
13.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
13.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
13.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
13.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
13.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
14 | Life Span for Machine: | 10 Years
15 | Maintenance:
15.1 |  | 2 Years Free Maintenance.
15.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
15.3 |  | Spare Parts list with code NO
16 | Training: | User /Nurses Training, by Specialist from the Supplier. And Biomedical Engineers
17 | Power supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
18 | Other Specification: | Please Specify Other Specification.
TMY-42 | Item Description | Hand Held Pulse Oximeter
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Type: | Hand Held Type
7 | SpO2 Range, %: | 70-100 Or Better.
7.1 | Accuracy, %: | ±3% Or Better.
8 | Pulse Rate bpm: | 30-240 Or Better.
8.1 | Accuracy: | ±5%.Or Better.
9 | Perfusion Index and/or Signal Strength Indicator. | Required, continuous monitors; preferred, spot check monitors.
10 | Response Time Sec: | Depends on application.
11 | Alarms :
11.1 | Audible and visual: | High/low SpO2 and pulse rate, sensor off, sensor failure, low battery
11.2 | Visual only: | High/low SpO2 and pulse rate, sensor off, low battery
12 | Alarm Override: | Preferred
12.1 | Reactivation method: | Automatic or manual
12.2 | Volume Control: | Required
13 | Self -Test Mode: | Required
14 | Sensor Types:
14.1 |  | Reusable Type
14.2 | Patient range: | Adult, pediatric, infant, neonate
14.3 | Cable length, m (ft): | 1.5 or More
15 | Battery:
15.1 | Type: | Rechargeable lithium-ion / or better.
15.2 | Life, hr: | 6- 12 Hr
15.3 | Recharge time, hr: | ≤3.
16 | Certification from the Manufacturer:
16.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
16.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
16.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
16.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
16.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
16.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
16.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
16.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
16.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
17 | Life Span for Machine: | 10 Years
18 | Maintenance:
18.1 |  | 2 years free Maintenance.
18.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
18.3 |  | Spare Parts list with code NO
19 | Required accessories:
19.1 |  | 1 X SpO2 Adult Resuable Sensor
19.2 |  | 1 X SpO2 Pediatric Resuable Sensor
19.3 |  | 1 X Charger
19.4 |  | 1 X Carying Bag
20 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
21 | Battery: | Rechargeable battery with built in charger with indicators for charging and low battery
22 | Other Specification: | Please Specify Other Specification.
TMY-43 | Item Description | Washing Machine 15 Kg Top Loader or Front Loader.
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
4 | Country Of origin: | Please specify Country of origin of the offered equipment
5 | Design & Quality: | Heavy Duty, High Quality
6 | Key Features:
6.1 |  | Take all the hassle out of doing laundry with cool and capable 15kg Top Loader or Front Loader Washing Machine. Enjoy the ease and convenience of our One Touch Smart Wash technology and get the perfect wash every time with every load.hnology and get the perfect wash every time with every load.
6.2 | Load Capacity: | 15kg
6.3 |  | One Touch Smart Wash Technology
6.4 | Power Consumption: | 600W
6.5 | Operating System: | Silent
6.6 | Alarms | Included
6.7 | Washing Programme: | Included
7 | Cover: | Transparent
8 | Body surface finish: | High -quality plastic body
9 | Power supply: | 220V - 50/60Hz
10 | Estimated life-Span of the Item: | At Least 7 Years
11 | Warranty: | Minimum of 1 Year.
12 | Other Specification: | Please Specify Other Specification.
TMY-44 | Item Description | Cupboard Metal Steel 2-Doors Storage
Qty |  | 3
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Type: | Cupboard metal Steel 2-Doors Storage
7 | Composition: | from one piece , contain of 2 doors , entirely metallic Equipped with 4 metallic adjustable shelves , Heavy- gauged metal framing with clip and bolted assembly , can status each one in different level .
8 | Construction: | Made from entirely metallic , from matter Heavy-gauged , Epoxy-Painted and oven-baked components
9 |  | the cupboard contain from 2 metal lockable doors with heavy duty hinges , to make easy to open and close doors and accuracy
10 |  | Equipped the doors key kind one upset , doorknob each other door designer and position form proportionate to make easy to open & close doors
11 | Height : | 200 cm at least
12 | Width : | 100 cm at least
13 | Depth : | 45 cm at least
14 | Life Span for Item: | 7 Years
15 | Warranty: | Minimum of 1 Year.
16 | Other Specification: | Please Specify Other Specification.
TMY-45 | Item Description | Window Curtain
Qty |  | 3
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Design & Quality: | Heavy Duty, High Quality
5 | Genaral Specification:
5.1 |  | Window Curtain
5.2 |  | It should not need to be maintained periodically, because it has a long life.
5.3 |  | It should be resistant to sunlight.
5.4 |  | It should be easy to install and use.
5.5 |  | Ease of cleaning, because it consists of Cloths Parts that are easy to clean.
5.6 |  | Provide an aesthetic view.
5.7 | Dimensions | 200 L X 150 H cm
6 | Life Span For Item: | 10 Years
7 | Warranty: | Minimum of 1 Year.
8 | Other Specification: | Please Specify Other Specification.
TMY-46 | Item Description | Admission Bed , 3 Section
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Design & Quality: | Bed, designed 3 Sections standard wards .
7 | Construction: | Epoxy-Coated main frame constructed from heavy duty section steel tube
7.1 |  | Removable Chrome-Plated steel bed ends fitted with 10mm laminated plastic panels
7.2 |  | 3 Sections mattress Platform fitted with Epoxy-Coated steel weldmesh
8 |  | Backrest and knee brake adjustable by hydraulic action
9 |  | The bed shall be mounted on 4x125mm (minimum) castors, 2 swivelling with simultaneous brakes
10 |  | movable bed sids,four five- inch castors two of them are totally lockable,adjustable
11 | Guiding Dimentions of bed: | 210-220L x 90-100Wx 30-60H cm
11.1 | Accessories:
11.2 |  | Three sections foam. Mattress, 10-12cm at least High density, Premium-grade cotton felt and high-density urethane foam are intertufted to eliminate component shifting and to increase mattress life with Durable vinyl cover, anti-bacterial, anti-static, acid-resistant and waterproof for easy use and care
11.3 |  | IV rod, 2 hooks with adjustable height
11.4 |  | Side folding-down rails (pair)
2 X Pillow With Cases
1 X Ad
12 | Locker, Bedside | Heavy duty, Flexible, high quality
12.1 | Construction: | Made from the same Material and same Colors of bed,better From the same Beds Manufacturer
12.2 |  | 4 castors (Orientable)
12.3 |  | Upper drawer, with nylon ball bearing sliders. Internal Plastic receiver
14 |  | Over Bed Table Heavy duty, Flexible &Foldable, high quality
14.1 | Over Bed Table Intgreted or Sparately: | Please specify
14.2 |  | Gas Spring Height Adjustment By Lever.
14.3 |  | Table Top Made Of Playwood.or equivalen Material for easy cleaning
14.4 | Height Adjustment Range : | 700-1200mm.
15 | Life Span for Item: | 7 Years
16 | Warranty: | Minimum of 2 Years For All.
17 | Other Specification: | Please Specify Other Specification.
TMY-47 | Item Description | Anesthesia Machine with Ventilator
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
6 | System Components:
6.1 |  | Original cart with antistatic braked castors, shelves and 2-3 drawers.The basic unit must form rigid stable and easy for use.
6.2 |  | Ventilator
6.3 |  | Multi vaporize capability
6.4 |  | Regulators and flow meters
6.5 |  | CO2 absorber
6.6 | Patient Type: | Adult, Pediatric, Infant
6.7 |  | Patient monitor
7 | Circulation System:
7.1 |  | Compact
7.2 |  | Sterilizable
7.3 |  | Changeable from rebreathing (closed) to non-rebreathing (open)
8 | Gas Mixing System:
8.1 |  | Mechanical safety system
8.2 |  | O2 flow meters (One tube)
8.3 |  | N2O flow meters (One tube)
8.4 |  | Air flow meters (One tube)
8.5 |  | Nitrous oxide blockage with constant alarm
8.6 |  | N2O , O2 ratio control
8.7 |  | Control of gas supply for O2 and air and for all gases
8.8 |  | Safety valve built in to patient system
9 | CO2 absorber:
9.1 |  | Bag /ventilator switch
9.2 |  | By-pass switch
9.3 |  | Pressure reading (Manometer or Electronic)
9.4 |  | canester ( 2 kg capacity or 1.5 L)
10 | Vaporizer: | Must be from ( Penlon Or Drager)
10.1 |  | Selecta tic type
10.2 |  | Two vaporizers
10.3 |  | Fast coupling system
10.4 |  | Supplied with sevoflurane
10.5 |  | Supplied with isoflurane
11 | O2/N2O Cylinders Yokes with Cylinders: | Included
12 | O2–flush: | Included
13 | Mechanical Anti Hypoxic Device (AHD): | Included
14 | Central Gas Supply: | O2/N2O/Air, medical quality 3-6 bar
| (Pipeline Gas Inlet): | Selector switch for O2/N2O and O2/Air
15 | Flow Meters:
15.1 |  | Color coded
15.2 |  | O2 : 0.05-10 L/min or better
15.3 |  | Air: 0.2-12 L/min or better
15.4 |  | N2O: 0.05-10 L/min or better
16 | Ventilator: | Electronically controlled
16.1 | Bellow system: | Range from pediatric to adults
Standing , easily seen
16.2 | Ventilation Modes (selectable):
16.2.1 |  | Stand by
16.2.2 |  | Spontaneous
16.2.3 |  | Manual
16.2.4 |  | CMV (Pressure control/ Volume control)
16.2.5 |  | SIMV
16.2.6 |  | pressure support, advanced modes
16.3 | Monitor Parameter Readings:
16.3.1 |  | Volume
16.3.2 |  | Pressure (Peak , Peep)
16.3.3 |  | Oxygen
16.3.4 |  | Frequency
16.3.5 |  | I : E ratio
16.4 | Monitor Graphic Waveforms: | Volume
Pressure
16.5 | Technical data :
16.5.1 | Respiration Frequency: | 4-60 bpm.
16.5.2 | Inspiration Flow: | 0-180 L/min
16.5.3 | Tidal Volume: | 20-1500 ml
16.5.4 | I : E Ratio: | Please Specify
16.5.5 | Pressure limit, cm H2O: | Adjustable
16.5.6 | PEEP: | 0-20 cm H2O or better
17 | Audible and Visual Alarms:
17.1 |  | Low/ high pressure
17.2 |  | Low O2 concentration
17.3 |  | Low/ high Volume
17.4 |  | O2 inlet supply failure
17.5 |  | Alarm interruption 60 sec. with rest function
17.6 |  | battery alarm
17.7 |  | Power failure alarm
18 | Accessories
18.1 |  | Anesthesia Masks Resuable (all sizes) , Adult, Pediatric, Infant.
18.2 |  | Sphygnomanometer
18.3 |  | Re-breathing bags Resuable (all sizes) , Adult, Pediatric, Infant.
18.4 |  | Re-breathing Circles Resuable (all sizes) , Adult, Pediatric, Infant.
18.5 |  | Magil circuit
18.6 |  | Jackson-read modification of Ayri’s T piece
18.7 |  | O2 sensore
18.8 |  | All needed accessories to insure full use for adults, pediatrics and neonat.
18.9 |  | Operation Manual
18.1 |  | Service Manual
18.11 |  | Spare parts list with code NO.
19 | Capnography/CO2 Monitoring: | Sidestream Or Mainstream
19.1 | Halothane: | 0 - 8.5 Vol % Or Better
19.2 | Isoflurane: | 0 - 8.5 Vol % Or Better
19.3 | Enflurane: | 0 - 10 Vol %
19.4 | Sevoflurane: | 0 - 10 Vol %
19.5 | Desflurane: | 0 - 20 Vol %
19.6 | Accuracy: | ± 0.2 Vol Or Less
20 | Self Diagnostic and Error / Calibration Message: | Included
21 | Power Supply: | 100 - 240 V, 50 / 60Hz
Internal Rechargeable Battery
22 | Essential Requirement:
22.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
22.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
22.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
22.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
22.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
22.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
22.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
22.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
22.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
22.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
22.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
23 | Estimated life-Span of the Device: | At Least 10 Years
24 | Maintenance:
24.1 |  | 2 Years Free Maintenance.
25 | Training: | User /Nurses training, by Specialist from the Supplier.
26 | Other Specification: | Please Specify Other Specification.
TMY-47-1 | Item Description | Patient Monitor of Anesthesia Machine
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Type: | Modular, use separate parameter modules or composite module with facility removing any parameter will not affect the others patient type: adult, pediatric, neonate High quality
Upgradeable by software and hardware
7 | Display size and Type: | 12-15 Inches Or Better. Touch Screen
High resolution multi-color display
Medical type (preferable) or supported with isolation transformer
8 | Displayed information: | Min. 6 vital waveforms, cab be colored separately
Numeric data for the measured vital parameters
Vital parameters (72) hrs trends
Vital parameters alarms audio visual
9 | Defibrillation Protection: | Available for the measured vital parameters
10 | Required vital parameters:
10.1 | ECG:
10.1-1 | Leads: | 3-5 Leads facility
10.1-2 | Values: | ECG lead waveform, label, HR gain as min.
10.1-3 | HR Range: | 30-200 bpm Or Better.
10.1-4 | Alarms: | Leads off, Hi + Low HR
10.1-5 | Gain: | 5, 10, 20 mm/mV Or Better.
10.1-6 | Preferable Items: | ST, Arrhythmia, cascade ECG
10.2 | Respiration:
10.2-1 | Technology: | Impedance
10.2-2 | Values: | Respiration waveform, R.R.
10.2-3 | R.R. Range: | 5-80 bpm Or Better.
10.2-4 | Alarms: | Hi & low R.R.
10.2-5 | Apnea Alarm: | 15-25 sec. Preferable
10.3 | NIBP:
10.3-1 | Technology: | Oscillometric
10.3-2 | Values: | SYS , DIA , MEAN , P.R.
10.3-3 | Modes: | Auto, Manual
10.3-4 | Cuff Pressure Range:
10.3-4-1 |  | Up to 250mmHg Adult
10.3-4-2 |  | Up to 200mmHg Pediatric
10.3-4-3 |  | Up to 150mmHg Neonate
10.3-4-4 | Alarms: | Hi and low, SYS, DIA
10.4 | SPO2:
10.4-1 | Values: | SPO2 waveform, SPO2% , P.R.
10.4-2 | SPO2 range | 50 - 100 %
10.4-3 | Pulse Rate Range: | 25-200ppm Or Better.
10.4-4 | Alarms: | Hi & low SPO2 + P.R. sensor off
10.5 | Temperature:
10.5-1 | Values: | T1 and or T2
10.5-2 | Range: | 30-45°C Or Better.
10.5-3 | Alarm: | Hi & low
10.6 | ETCO2:
10.6-1 |  | Should be Main Stream capnography with display of CO2 and digital Values of EtCO2, FiCO2 & RR.
10.6-2 | EtCO2 Range : | 0-99 mmHg.
10.6-3 | FiCO2 : | 0 to 20 mmHg.
10.6-4 | Flow rate : | 50ml/min.
10.7 | Invasive BP:
10.7-1 |  | Should have 2 channel Invasive Blood pressure (IBP) measurement.
10.7-2 |  | Should have waveform IBP1 and IBP2.
11 | Required Accessories:
11.1 |  | 1 X 3/5 Leads wire ECG cable Adult.
11.2 |  | 1 X 3/5 Leads wire ECG cable Pediatric.
11.3 |  | 1 X NIBP reusable cuff large Adult
11.4 |  | 1 X NIBP reusable cuff pediatric
11.5 |  | 1 X NIBP reusable cuff Infant
11.6 |  | 1 X NIBP connection hoses
11.7 |  | 1 X SPO2 reusable finger probe + Extension cable adult.
11.8 |  | 1 X Temperature reusable sensor skin type for , pediatric and Infant
11.9 |  | 1 X EtCO2 module with all accessories. ( Cable-Sensor)
11.1 |  | 2 X IBP Reusable Transducers with cable
11.11 |  | 10 X IBP Disposable pressure transducers
11.12 |  | Operation Manual
11.13 |  | Service Manual
11.14 |  | Spare parts list with code NO.
12 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
12.1 | Internal back-up battery | Operating time, hr : ≥2
13 | Essential Requirement:
13.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
13.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
13.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
13.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
13.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
13.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
13.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
13.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
13.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
13.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
13.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
14 | Estimated life-Span of the Device: | At Least 10 Years
15 | Maintenance: | 2 Years Free Maintenance.
16 | Training: | User /Nurses training, by Specialist from the Supplier.
17 | Other Specification: | Please Specify Other Specification.
TMY-48 | Item Description | Operating Table "Major''
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Operating Table type: | Electric Or Hydrolic
7 | Maximum weight capacity, kg: | Static: 200-250
Articulated: ≥220
8 | Table Positions: | Specify
9 | Degrees From Horizontal: | Specify
| Trendelenburg: | ≥35
10 | Reverse Trendelenburg: | ≥30
11 | Lateral tilt: | ≥25
| Vertical range, cm: | 60-100 height-adjustable
12 | Table Section: | ≥5
13 | Number of Core Sections: | ≥3
| Degrees from Horizontal: | Specify
14 | Back Section: | 70 to30
15 | Foot/leg Section: | 20 to 90
| Head Section: | 45to85
16 | Controls: | Hand or foot and Remote Control
17 | Siderail: | Including.
| C-arm accessible: | Including.
18 | Base attachment: | Including.
19 | Column Housing: | Including.
| Conductive Casters: | Including.
20 | Caster Lock: | Including.
21 | Certification from the Manufacturer:
21.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
21.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
21.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
21.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
21.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
21.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
21.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
21.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
21.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
22 | Life Span for Machine: | 7 Years
22.1 | Maintenance:
22.2 |  | 2 Years Free Maintenace.
22.3 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
22.4 |  | Application Software and Interface Connection Included.
22.5 |  | Spare Parts list with code NO
23 | Accessories:
23.1 |  | Anesthesia screen L type with clamp 1 piece
23.2 |  | Leg Support with clamp 2 pieces
23.3 |  | Arm Support with clamp 2 pieces
23.4 |  | Body Support 1 piece
23.5 |  | IV rod, 1 set, adjustable height with clamp
23.6 |  | Shoulder rest / support with pad, 2 pieces
23.7 |  | Head Plate 1 piece
23.8 |  | Chest & waist support
24 | Other Specification: | Please Specify Other Specification.
TMY-49 | Item Description | Electro Surgical Unit 400 W
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Generator Type: | Please Specify.
7 | Frequency: | High Frequency.
8 | Modes:
8.1 | Monopolar Mode: | Cut - Coagulate
8.1-1 | Cut Mode: | Low cut: maximum power of 100-150 w, Adjust output power Or Better. Pure cut: maximum power of 300-400w, Adjust output power Or Better. Blend mode: maximum power of 100-300w, Adjust output power Or Better.
8.1-2 | Coag Mode: | Maximum 100-200 watts Or Better. Please specify The Modes.
8.2 | Bipolar mode: | Please specify The Modes.
8.3 | Self-test Mode: | Including.
8.4 | Power Setting Displayed: | Including.
8.5 | Cooling : Convection | Including.
9 | Alarms: | Visual - audible
9.1 | For Plate Continuity: | Including.
9.2 | Excess Power: | Including.
9.3 | Internal Error Alarm: | Including.
9.4 | Plate voltage alarm: | Including.
9.5 | Patient Electrode Safety System: | Including.
9.6 | Defeatable-Inaudible Volume: | Including.
10 | Certification from the Manufacturer:
10.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
10.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
10.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
10.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
10.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
10.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
10.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
10.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
10.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
11 | Life Span for Machine: | 10 Years
12 | Maintenance:
12.1 |  | 2 Years Free Maintenace.
12.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
12.3 |  | Application Software and Interface Connection Included.
12.4 |  | Spare Parts list with code No
13 | Accessories | Supplied with :
13.1 | Monopolar Footswitch: | Supply 1 pice
13.2 | Bipolar Footswitch: | Supply 1 pice
13.3 | Reusable hand switching Pencil: | Supply 2 pice
13.4 | Reusable Patient Plate: | Supply 2 pice
13.5 | Bipolar Forceps: | Supply 2 pice
13.6 | Forceps Cord: | Supply 2 pice
13.7 | Universal Adaptor: | Supply 1 pice
13.8 | Trolley: | Included Original Type
14 | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
15 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
16 | Other Specification: | Please Specify Other Specification.
TMY-50 | Item Description | Laryngoscope
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality,
6 | Laryngoscope:
6.1 | Genaral Specification:
6.2 |  | Includes a Cylindrical Handle, Slightly Ribbed, Single Piece, Ø > 25 mm The Handle Allows Fitting of Interchangeable Blades of Different Sizes. The Handle Contains a Robust On/Off Switch.
6.3 |  | Handle is Made from Non-Ferrous Material and Sealed Against Ingress of Liquids.
6.4 |  | Equipped with LED Technology Light Source.
6.5 |  | Uses Fibre Optic for Transmission of the Light to the Tip of the Blade.
6.6 |  | Stud Contact, Fitting Various Sizes and Types of Depressors
6.7 |  | Blades Are Able to Withstand Autoclaving.
6.8 |  | With a Set of Four Blades Made from Stainless Steel:
6.9 |  | 1 X MacIntosh type: Curved Nr 2, length 110mm ± 5 mm.
6.1 |  | 1 X MacIntosh type: Curved Nr 3, length 135mm ± 5 mm.
6.11 |  | 1 X MacIntosh type: Curved Nr 4, length 155mm ± 5 mm.
6.12 |  | 1 X Miller type: Straight Nr 1, length 100mm ± 5 mm.
6.13 | Accessories:
6.13.1 |  | 1 X Carying Bag.
6.13.2 |  | 2 X Alkaline Batteries.
6.13.3 |  | 4 X Spare LED.
17 | Life Span for Item: | 5 Years
18 | Warranty | Minimum of 1 Year.
19 | Other specification | Please Specify Other Specification.
TMY-51 | Item Description | X-Ray Viewer
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality.
6 | Type: | Wall mounted X-Ray film viewer
7 | Construction: | made from acrylic translucent diffusing materials . screen thickness ≥ 3mm approx. automatic gripping mechanism included
8 | Functionality: | For general radiology applications. Appropriate for viewing a board band rang of film sizes
9 | Viewing panel size (L x W) , cm: | (35 x 43)
10 | Light: | LED
11 | Number of Lamps: | Please specify
12 | On – Off Switch: | Included
13 | Power Supply: | 220 V / 50 Hz
14 | Warranty: | Minimum of 1 Year.
15 | Other Specification: | Please Specify Other Specification.
TMY-52 | Item Description | Disinfection Ultraviolet Lamp With Infrared Sensor
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality,
5 | Type:
5.1 |  | Bacterial Lamp for Air Sterilization
5.2 |  | UVC shall be emitted thru vacuum ionizing tube (Weak permeation). Useful power shall be at least 45% of nominal power (30W)
6 | Lifetime: | Lifetime shall be at least 7500 Hrs. Radiation shall be witnessed by means of blue light emission
7 | Timer: | Timer shall assert the limit period of the ionized tube running
8 | Castors: | Four casters ,two with brakes
9 | Accessories: | Remote Control
9.1 |  | 4 Lamps as Spare Part.
10 | Power supply: | 100 to 240 V ~ ±10%, 50 Hz ( power cable Compatible with the Hospital electric outlet, plug ), Electrical Safety class 1,with indicators for power
11 | Estimated life-Span of the Device: | At Least 7 Years
12 | Maintenance: | 2 Years free maintenance.
13 | Training: | User /Nurses training, by Specialist from the Supplier.
14 | Other Specification: | Please specify.
TMY-53 | Item Description | Split Unit Air Conditioner 2 Ton
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must have Lattest Version
5.3 | Quality: | High Quality,
6 | Cooling Capacity : | 5 kW (2 TR)
7 | Star rating : | 3
8 | Max. Ambient temp: | 45 ºC.
9 | Energy efficient ratio (EER) : | 3.2 Or Better.
10 | Electrical input: | 230V/50Hz/Single Phase
11 | Noise level Indoor unit: | ≤ 40 dB
12 | Air flow rate: | 425 CFM or more
13 | Fan speed: | 3 Steps.
14 | Function modes: | Auto/Cool/Fan/Dry – shall have power saving mode, temperature setting option should be available.
15 | Body surface finish: | Powder coated/high quality paint finish.
16 | Air filtering unit: | Dust proof and anti-bacteria filter
17 | Length of tubing: | 5 m or as per the installation requirement.
18 | Remote control: | Wireless with LCD display.
19 | Refrigerant type: | R-22 or R-410A or any other refrigerant as per current industry standard.
20 | Estimated life-Span of the Item: | At Least 10 Years
21 | Warranty: | Minimum 1 Year.
22 | Other Specification: | Please Specify Other Specification
TMY-54 | Item Description | Hygrometer Digital
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed.
5.2 |  | The Machine should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
6 |  | Fast Response time
7 |  | High Accuracy
8 |  | Simultaneous Indication of Temperature and Humidity
9 |  | Automatic Stiwch-off
10 |  | Calculation of Dew Point and Wet-Bulb Temperature
11 |  | Background Lights
12 |  | Data Memory of max. value
13 |  | Compact and Light
14 | Technical specifications:
14.1 | Measurement Ranges: | 0 to 100% r.h , -30 to 100°C
14.2 | Resolution: | 0.01% r.h , 0.01°C
14.3 | Accuracy: | ±2.0% r.h. at 25°C , ±0.5°C at 25°C
14.4 | Display : | LCD Type
14.5 | Battery: | 9V Battery (PP3) Or Equivalent
14.6 | Response Time % r.h: | About 10 Sec Or Better.
15 | Accessories:
15.1 |  | 1 X Carrying Case
15.2 |  | 1 X Spare Part Battery.
16 | Estimated life-Span of the Item: | At Least 7 Years
17 | Warranty: | Minimum of 1 Year.
18 | Other Specification: | Please Specify Other Specification
TMY-55 | Item Description | Stainless Steel Table Work
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed.
5.2 | Quality: | High Quality
6 |  | Stainless Steel Table Work
7 | Material: | Stainless Steel
8 | Surface Finish: | Polish
9 | Load Capacity: | 150KG/tier
10 | Color: | Silver
11 | Dimensions: | 1800x800
11.1 | Length: | 180-200 cm
11.2 | Width: | 70-80 cm
11.3 | High: | 80-90 cm
12 | Estimated life-Span of the Item: | At Least 7 Years
13 | Warranty: | Minimum of 1 Year.
14 | Other Specification: | Please Specify Other Specification
TMY-56 | Item Description | Mayo Table Stainless Steel
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must have lattest Version
5.3 | Quality: | High Quality,Safety, Ease of Use
6 | Material of Construction: | Stainless Steel 304 or better.
7 | Shelves: | One Bin.
8 | Dimensions:
8.1 |  | Depth = 45 Cm Approx.
8.2 |  | Width = 60 Cm Approx.
8.3 |  | Height = 80 Cm Approx.
9 |  | Trolley for Instruments.
10 | Castors: | Mobile on 4 antistatic castors, 2 of which shall have foot brakes
11 |  | The unit shall be height adjustable with rebatted edges
12 |  | Equipped with stainless steel bin (At least 15L Capacity ) and Bowls' rack
13 | Accessories & Options: | Please Specify
14 | Estimated life-Span of the Item: | At Least 7 Years
15 | Warranty: | Minimum of 2 Years.
16 | Other Specification: | Please Specify Other Specification
TMY-57 | Item Description | Instrument Trolley
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material of Construction: | Heavy Duty Stainless Steel 304 or Better and High Quality
7 | Shelves: | Two Shelves in the Top and Bottom , Edge Around the Shelves .
8 | Dimensions:
8.1 | Length : | 80-90 Cm Approx.
8.2 | Width : | 50-60 Cm Approx.
8.3 | height : | 75-85 Cm Approx.
9 | Handle Bar: | Included ( on Short Sides of the Trolley).
10 | Castors: | Four Casters ,Two with Brakes
11 | Accessories & Options: | Please Specify
12 | Life Span for Item: | 7 Years
13 | Warranty: | Minimum of 2 Years.
14 | Other Specification: | Please Specify Other Specification
TMY-58 | Item Description | Bench Top Steam Sterilizer With Drayer
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
6 | Air Removal Method: | To Facility Requirements
7 | Chamber Material: | 314 Stainless Steel
7.1 | Type: | Class B
7.2 | Number of trays: | 2-3.
8 | Water Reservoir Capacity , L: | 18-24 L
9 | Preprogrammed Cycles, Type: | Wrapped, Unwrapped
10 | Cycle Number ,Type: | Please Specify
11 | Minimum Cycle Time ,Min: | Please Specify
12 | Automatic Cycle Shut off: | Included.
13 | Temperature Range , °C : | 100-134 °C
13.1 | Overheat shutoff: | Included.
14 | Pressure Purge: | Please Specify
15 | Controller Type: | Please Specify
15.1 | Programmable: | Included.
16 | Alarm Conditions:
16.1 | Indicator type: | Audible, visual
16.2 | Chamber pressure or temperature Failure: | Audible, visual
16.3 | Component Failure: | Audible, visual
16.4 | Cycle disruption or Failure: | Audible, visual
17 | Data Management:
17.1 | Recorder: | Included.
17.2 | Printer: | Included.
18 | Power Supply, VAC: | 110-230 50/60 Hz
18.1 | Consumption, Kw: | Please Specify
19 | Drayer Function: | Included. Vaccum Type
20 | Accessories:
20.1 |  | Steel container : plain basket with handles.
20.2 |  | Paper for printer Qty: 10 rolls
20.3 |  | Spare Part Heater 1 Pcs.
20.4 |  | Rubber Gasket for Door 1 Pcs.
21 | Certification from the Manufacturer:
21.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
21.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
21.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
21.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
21.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
21.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
21.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
21.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
21.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
22 | Life Span for Machine: | 10 Years
23 | Maintenance:
23.1 |  | 2 years free maintenace.
23.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
23.3 |  | Application Software and Interface Connection Included.
24 | Power Supply: | AC 110-240V, 50/60Hz single-phase electrical source.
24.1 |  | Compliant with IEC 60601.
24.2 |  | Built-in protections against over-voltage and over-current line conditions.
25 | Other Specification: | Please Specify Other Specification
TMY-59 | Item Description | Double Sink Wash Basin Stainless Steel
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed.
5.2 | Quality: | High Quality
6 | Material: | 304 Grade Stainless Steel.
7 | Internal Bowl Size of : | 24” Long x 26” Wide x 15” Deep for Every One.
8 | Thickness: | 1.5 mm
9 |  | Splash back to the rear.
10 |  | Mounted on Stainless Steel Adjustable Legs.
11 | Estimated life-Span of the Item: | At Least 7 Years
12 | Warranty: | Minimum of 1 Year.
13 | Other Specification: | Please Specify Other Specification
TMY-60 | Item Description | Emergency Cart. (Crash Cart with Complete Access. )
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Item Must Have Lattest Version
5.3 | Quality: | High Quality,
6 | General Description: | Emergency Cart, with accessories.
7 | Technical specifications: | Emergency trolley or (Emergency Crash Cart) is a set of trays/drawers/shelves on wheels used in hospitals for transportation and dispensing of emergency medication/equipment at site of medical/surgical emergency for life support protocols potentially to save a patient’s life.
8 | Constructions: | Stainless steel top and shelf & equipped with 4 - 6 removable coloured bins made of moulded plastic.
9 | System Configuration: | Emergency trolley (Emergency Crash Cart) with removable coloured bins, storage units, fitted with oxygen cylinder holder and electric lamp holder and four swivels castors.
10 | Dimensions: approx. | 960mm L x 500mm W x 1545mm H.
11 |  | Lockable storage units – 3 drawers (stainless steel or moulded plastic). Wood or wood laminate construction drawers are NOT acceptable.
12 |  | To be fitted with stainless steel, height adjustable, twin hook/loop, IV pole assembly
13 | Castors: | Fully, 360 deg. swivel castors/wheels, size 125mm dia with at least one castor/wheel to .have locking/brake mechanism.
14 | Shelves: | SS shelves 14 G thick stainless steel.
15 |  | Top shelf to have stainless steel guard rail approx.35mm above surface.
16 |  | Fitted with epoxy powder coated oxygen cylinder holder and electric lamp holder with clamp and cardiac massage board.
17 |  | Must be capable of carrying ECG Monitor/defibrillator and a suction apparatus.
18 | Accessories, spares and consumables: | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
19 | Estimated life-Span of the Item: | At Least 7 Years
20 | Warranty: | Minimum of 2 Years.
21 | Other Specification: | Please Specify Other Specification
TMY-61 | Item Description | Infant Warmer
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Type: | Mobile on Swivel Castors , 2 (locks).
7 | Control Unit:
7.1 | Heat Control, Type: | Automatic or manual
7.1.1 | Range, °C: | 34-37 °C Or Better
7.2 | Heater indicator: | Required
8 | Temperature:
8.1 | Set-point Display: | Please specify
8.2 | Skin-temp Display: | Please specify
9 | Alarms:
9.1 | Temperature:
9.1.1 | Tone: | Variable or fixed
9.1.2 | Type: | Audible and visual
9.1.3 | Trigger from setpoint, ±°C/°F | 0.5-1 Or Better.
9.2 | Manual Mode: | Continuous until corrected, automatic shut off if not fixed within 10 min
9.3 | Sensor Disconnect: | Audible and/or visual
9.4 | Power Failure: | Audible and/or visual
10 | Bassinet: | Required
11 | Radiant Heat Source: | Ceramic
12 | Examination Light: | LED
13 | X-ray tray: | Required
14 | Drawers/shelves: | Required
15 | Timer: | Required
16 | Mattress: | Required
17 | Accessories:
17.1 |  | 1 X IV Pole.
17.2 |  | 2 X Spare Heaters
17.3 |  | 2 X Oxygen Holder.
17.4 |  | 3 X Fuses.
17.5 |  | 2 X Spare Led Bulb
17.6 |  | 1 x Operation Manual
17.7 |  | 1 x Service Manual
18 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
19 | Essential Requirement:
19.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
19.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
19.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
19.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
19.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
19.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
19.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
19.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
19.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
19.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
19.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
20 | Estimated life-Span of the Device: | At Least 10 Years
21 | Maintenance: | 2 Years free maintenance.
22 | Training: | User /Nurses training, by Specialist from the Supplier.
23 | Other Specification: | Please specify.
TMY-62 | Item Description | Fire Extinguisher 10Kg Dry Powder
Qty |  | 3
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
5 | Design & Quality: | Heavy Duty, High Quality
6 | Capacity: | 10 Kg
7 | Fire performance : | 43A 233B C
8 | Extinguishing Agent: | ABC 40% Dry Powder
9 | Operating temperature: - | 30°C to 60°C
10 | Average Discharge Time: | 18.47sec
11 | Test Pressure: | 26bar
12 | Max Allowable Pressure: | 18.00bar
13 | Volume of Equipment : | 11.90Lt
14 | Cylinder Material: | St 12
15 | Cylinder Height: | 500 ± 5 mm
16 | Cylinder Diameter: | 200 ± 1.5 mm
17 | Cylinder Neck Thread: | 30 x 1,5 mm
18 | Valve Material: | HPb59-1
19 | Safety Valve: | 22-26 bar
20 | Paint Finish : | RAL 3000
21 | Approx. Gross Weight: | 11.50Kg
22 | Approx. Dimensions: | 650 x 200 mm
23 | Work Pressure : | At 20°C 15bar
24 | Approx. Discharge Distance : | 5-10m
25 | Life Span For Item: | 10 Years
26 | Warranty: | Minimum of 1 Year.
27 | Other Specification: | Please Specify Other Specification
TMY-63 | Item Description | Infusion Pump W/Holder
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Mode:
6.1 |  | Macro : 1-999 ml/hr in 1ml./hr increment Or Better.
6.2 |  | Micro : 0.1-99.9 ml/hr in 0.1 ml/hr increment Or Better.
7 | Volume to be infused: | 1.0 – 999.9 ml or No limit Or Better.
8 | Infusion Set Used with the System: | should (open system) the user can use any infusion set
9 | Infusion Accuracy:
9.1 |  | Presettable rate
9.2 |  | At Least ±2% Or Better.
10 | Secondary Function:
10.1 |  | Free flow protection
10.2 |  | Front panel key pad lock
10.3 |  | Built-in self test and diagnostic
10.4 |  | System and malfunction alarms with audible and visual alarm indicators
10.5 |  | Silence or reset of audio alarms
10.6 |  | Air line detect system
11 | Message Display for Infusion Status: | Alarms, error messages etc.
12 | Automatic Flow:
12.1 |  | Rate calculation
12.2 |  | Drug calculation & drug dose programs
12.3 |  | Multi-dose capability
12.4 |  | Rate-volume and volume timer programming
13 | Accessories:
13.1 |  | 1 X Holder
13.2 |  | 1 X Operation Manual
13.3 |  | 1 X Service Manual
13.4 |  | 1 X Set of Spare Infusion Set For Adult & Pediatric. (each) 500 Set
14 | Power S:upply: | 100 to 240 V ~ ±10%, 50/60 Hz .
14.1 | Battery Pickup: | 4-6 Hrs.
15 | Essential Requirement:
15.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
15.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
15.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
15.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
15.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
15.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
15.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
15.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
15.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
15.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
15.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
16 | Estimated life-Span of the Device: | At Least 10 Years
17 | Maintenance: | 2 Years free maintenance.
18 | Training: | User /Nurses training, by Specialist from the Supplier.
19 | Other Specification: | Please Specify Other Specification
TMY-64 | Item Description | Syring Pump W/Holder
Qty |  | 4
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Syring Pump: | Permits the use of standard disposable Syringes (sizes 10 – 50/60 ml) and extension sets, with auto syringe size sensing.
7 | Infusion Volume: | 0.1 to 999 ml in 0.1ml increments in lower ranges Or Better.
8 | Infusion rate adjustable: | from 0.1 – 99.9 ml/hr in (0.1ml/hr steps) Or Better.
9 | System and malfunction: | alarms with visual and audible alarm indicators
10 | Display:
10.1 |  | Digital display or better of volume infused. Infusion rate, alarm prompts etc.
10.2 |  | Purge/ bolus facility
10.3 |  | Volumetric accuracy At Least ± 2%. Or Better.
10.4 |  | Occlusion Pressure limit (from 0 to 750 mmHg)
10.5 |  | Built-in self test and diagnostics
11 |  | DPS (dynamic pressure sensor) Included
12 |  | Special feature of the product (other than mentioned on list)
13 |  | Accessories if any (to put machine/ system into operation)
14 |  | Consumable if any (to put machine/ system into operation)
15 | Accessories:
15.1 |  | 1 X Holder
15.2 |  | 1 X Operation Manual
15.3 |  | 1 X Service Manual
16 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
16.1 | Battery: | 4-6 Hrs.
17 | Essential Requirement:
17.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
17.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
17.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
17.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
17.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
17.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
17.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
17.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
17.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
17.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
17.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
18 | Estimated life-Span of the Device: | At Least 10 Years
19 | Maintenance: | 2 Years Free Maintenance.
20 | Training: | User /Nurses training, by Specialist from the Supplier.
21 | Other Specification: | Please Specify Other Specification
TMY-65 | Item Description | Appendicectomy Set
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material: | 304 Stainless Seel Or Better
7 | Each Set Contents of:
7.1 |  | 2 X Gallipot/Cup S/S 180 ml
7.2 |  | 2 X Tray Kidney shape without lid 300 X 50mm Deep.
7.3 |  | 1 X Tray Instrument with lid 310X 200 X63 mm deep
7.4 |  | 3 X Forceps Artery Halsted Mosquito COF BJ 120mm. S/S
7.5 |  | 3 X Forceps Artery Halsted Mosquito straight BJ 120mm. S/S
7.6 |  | 3 X Forceps Artery Spencer Wells CO Flat BJ.S/S 200mm
7.7 |  | 3 X Forceps Artery Spencer Wells Straight BJ.S/S 200mm
7.8 |  | 1 X Forceps Dissecting Bonney's 1X 2 Teeth S/S 150 mm
7.9 |  | 1 X Forceps Dissecting, narrow, serrated jaws 200 mm
7.1 |  | 2 X Forceps Tissue Allis fine 4 x 5 Teeth BJ.S/S 150 mm
7.11 |  | 4 X Forceps Sponge Holding Rampley BJ S/S 240 mm
7.12 |  | 1 X Handle size 3 for Scalple blades size 10-15
7.13 |  | 1 X Handle size 4 for Scalple blades size 20-24
7.14 |  | 1 X Needel Holder Mayo Hegar BJ.S/S 180mm
7.15 |  | 1 X Scissors Iris sharp points Straight S/S SJ.115 mm
7.16 |  | 1 X Scissors Mc Indoe/Metzenbaum CO Flat SJ.S/S 190mm
7.17 |  | 1 X Retractor Deaver Medium Blade 25 mm Length
7.18 |  | 1 X Retractor Deaver Large Blade 51 mm Length 305 mm
7.19 |  | 2 X Retractor Langenback H/H 6x22 Blade, 210 mm long
7.2 |  | 1 X Needle Aneurysm S/S . Small 146 mm
7.21 |  | 2 X Pin Safety Mayo, forceps holding, 120mm
7.22 |  | 1 X Suction Tube Pool's S/S 225mm
7.23 |  | 1 X Diathermy Electrode
7.24 |  | 1 X Sterilizing and storing case St. St. with filters 230x130x50 mm
8 | Life Span for Item: | 10 Years
9 | Warranty: | Minimum of 1 Year.
10 | Other Specification: | Please Specify Other Specifications.
TMY-66 | Item Description | Ceasarian Surgical Set
Qty |  | 2
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material: | 304 Stainless Seel Or Better
7 | Each Set Contents of:
7.1 |  | 2 X Tray Kidney shape with Lid S/S 200 X 30 mm
7.2 |  | 2 X Tray Instrument with Lid 225X 125 X 50 mm Deep
7.3 |  | 12 X Forceps Artery Spencer Wells Straight BJ.S/S 125mm
7.4 |  | 2 X Forceps Artery Spencer Wells Straight BJ.S/S 200mm
7.5 |  | 2 X Forceps Dissecting Plain serrated jaws S/S 150 mm
7.6 |  | 2 X Forceps Dissecting Treves, 1x2 Teeth 150 mm
7.7 |  | 2 X Forceps Tissue Allis fine 4 x 5 Teeth BJ.S/S 150 mm
7.8 |  | 3 X Forceps Tissue Little wood 2 x 3 Teeth (bj).S/S 190 mm
7.9 |  | 6 X Forceps Towel Holding cross Action S/S 100mm
7.1 |  | 4 X Forceps Sponge Holding Rampley BJ S/S 240 mm
7.11 |  | 1 X Handle size 3 for Scalple blades size 10-15
7.12 |  | 1 X Handle size 4 for Scalple blades size 20-24
7.13 |  | 2 X Needle Holder Mayo Hegar BJ.S/S 180mm
7.14 |  | 1 X Scissors Mayo Straight SJ. S/S 200mm.
7.15 |  | 1 X Scissors Mc Indoe/Metzenbaum CO Flat SJ.S/S 190mm
7.16 |  | 1 X Retr. Doyen Ring Handle blade 95 mm Length 280 mm
7.17 |  | 1 X Retr. Lange beck H/H 13x44 Blade, 210 mm Long
7.18 |  | 1 X Retr. Morris Double-ended (52x38)&(52x64) mm
7.19 |  | 3 X Pin Safety Mayo, forceps holding, 120mm
7.2 |  | 1 X Suction Tube Pool's S/S 225mm
7.21 |  | 6 X Forceps Haemostasis Green Armitage 200 mm
7.22 |  | 2 X Sterilizing and storing case St. St. with filters 260x150x50 mm
7.23 |  | 2 X Gallipot/cup S/S 180 ml
8 | Life Span for Item: | 10 Years
9 | Warranty: | Minimum of 1 Year.
10 | Other Specification: | Please Specify Other Specifications.
TMY-67 | Item Description | General Surgical Set Large
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | Designed: | Heavy Duty Designed, New Model.
5.1 | Quality: | High Quality.
6 | Material: | 304 Stainless Seel Or Better
7 | Each Set Contents of :
7.1 |  | 2 X Bowl Lotion 250 mm
7.2 |  | 2 X Galipot/Cup S/S 180 ml
7.3 |  | 2 X Galipot/Cup S/S 250 ml
7.4 |  | 2 X Tray Kidney shape with Lid S/S 200 X 30 mm
7.5 |  | 2 X Tray Kidney shape with Lid S/S 300 X 50 mm
7.6 |  | 20 X Forceps artery Kocher COF BJ.S/S 180mm.
7.7 |  | 15 X Forceps Artery Spencer Wells CO Flat (bj).S/S 200 mm
7.8 |  | 6 X Forceps Dissecting Plain serrated jaws S/S 150 mm
7.9 |  | 1 X Forceps Dissecting Plain serrated jaws S/S 250 mm
7.1 |  | 1 X Forceps Dissecting narrow, serrated jaws S/S 200 mm
7.11 |  | 3 X Forceps Dissecting Lane's 2x3 teeth 180 mm
7.12 |  | 1 X Forceps Dissecting Waugh, fine serr.jaws 1x2 Tth 200 mm
7.13 |  | 4 X Forceps Tissue Allis fine 4 x 5 Teeth BJ.S/S 150 mm
7.14 |  | 8 X Forceps Towel Holding cross Action S/S 100mm
7.15 |  | 4 X Forceps Towel Moynihan, Double Teeth (bj) 190mm
7.16 |  | 5 X Forceps Sponge Holding Rampley BJ S/S 240 mm
7.17 |  | 1 X Handle size 3 for Scalple blades size 10-15
7.18 |  | 1 X Handle size 4 for Scalple blades size 20-24
7.19 |  | 1 X Needle Holder Kilner's 165 mm
7.2 |  | 2 X Needle Holder Mayo Hegar BJ.S/S 180mm
7.21 |  | 1 X Scissors Blunt Point Straight SJ. S/S 180mm
7.22 |  | 1 X Scissors Mayo’s CO Flat 200mm S/S
7.23 |  | 3 X Scissors Mayo’s Straight 200mm S/S
7.24 |  | 1 X Scissors McIndoe's CO Flat 190mm S/S
7.25 |  | 1 X Retr.Deaver #1-Blade 25mm, Length 305mm
7.26 |  | 1 X Retr.Deaver #5-Blade 75mm, Length 305mm
7.27 |  | 2 X Retr. Lange beck, H/H 26x64 Blade, 210 mm Long
7.28 |  | 2 X Retr. Morris Double-ended (52x38)&(52x64) mm
7.29 |  | 1 X Retr. Walton, malleable copper 13mm length 305mm
7.3 |  | 1 X Needle Aneurysm S/S . Small 146 mm
7.31 |  | 4 X Pin Safety Mayo, forceps holding, 120mm
7.32 |  | 1 X Suction Tube McIndoe
7.33 |  | 1 XSuction Tube Poole's S/S 225mm
7.34 |  | 2 X Sterilizing and storing case St. St. with filters 260x150x50 mm
7.35 |  | 1 X Wire Cutting, Tungsten .Carbide inserts 180mm.
8 | Life Span for Item: | 10 Years
9 | Warranty: | Minimum of 1 Year.
10 | Other Specification: | Please Specify Other Specifications.
TMY-68 | Item Description | Semi Automated Biochemistry Analyzer
Qty |  | 1
No. | Technical Specifications | Min MOH Requirement | Bidders Offer | Notes
1 | Manufacturer : | Please Specify Manufacturer Name.
2 | Model Number: | Please Specify Model Number of The Offered Equipment.
3 | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan, USA and UK
5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy Duty Designed, New Model.
5.2 |  | The Machine shuld be of Latest Model and must Have Lattest Technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 | Technical Specification : | Should be microprocessor controlled general purpose bi-chromatic photometer
6.1 | Filters: | At least 7 Filters.
6.2 | Ranging nm: | 340 to 700 nm Or Better.
6.3 | Light source : | Tungsten/ halogen or higher grade (led lamp)
6.4 | Methods: | End point, kinetic and two point kinetic measurement modes or more
7 | Printer: | Build-in
8 | Measurement Range : | from 0.001 to 2.300Abs or better
9 | Display: | LCD display. Touch Screen
10 | Control: | Quality control – two control/test QC survey of at least 30 points, Levy Jenny plot.
11 | Bandwidth nm: | 5 nm or lesser.
12 | Programme Memory: | Should have a test programme memory of 50 or more.
13 |  | Should be provided with sample carry over prevention facility.
14 |  | Aspiration should be based on Bellow/Peristaltic Pump/ Vacuum pump.
15 | Reagents : | Should provide: 500 ml of reagents for urea, S. creatine, S. bilirubin, sugar, cholesterol, and Quality control 5ml one each for normal, abnormal.
16 | Certification from the Manufacturer:
16.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
16.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
16.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
16.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
16.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
16.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
16.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
16.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
16.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
17 | Life Span for Machine: | 10 Years
18 | Maintenance:
18.1 |  | 2 Years Free Maintenace.
18.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
18.3 |  | Application Software and Interface Connection Included.
18.4 |  | Spare Parts list with code NO
19 | Accessories:
19.1 | bulb: | One Spare Bulb
19.2 | Fues: | 3 Fueses
20 | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
21 | Power Supplay: | 100 to 240 V ~ ±10%, 50/60 Hz Single Phase automatic range selection ( power cable Compatible with the Hospital electric outlet plug, 5 mt ), Electrical Safety class 1
22 | Other Specification: | Please Specify Other Specification.