## Specification
ADRA YEMEN ADRA-YEMA II -0274-2024
ITEM No 1 | Gynecological delivery bed ( Hydraulic Type)
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The Item Must have lattest Version
5.3 | Quality: | High Quality,
5.4 | Supplier: | the supplier is required to be a certified distributor or agent of the company.
5 | Type: | Hydraulic Type
6 |  | For Delivery & Obstetric Use
7 | Materials: | Three sections,Made of steel, Epoxy-Painted with sturdy sampling of material .or any Equivalent rust-proof Material
7.1 |  | heavy duty ABS mattress base, 80 mm durable hygienic mattress,
8 |  | collapsible side rails, leg section mattress, adjustable foot supports, retractable foot section
9 | Adjustments:
9.1 | Height: | from 550 to 900 mm
9.2 | Back section: | 70 degree gas spring assisted
9.3 | Trendlenburg: | (+ 12 degree & -3 degree)
10 |  | With the adition of Pelvic position
11 | Permissible Load: | 120 - 150 Kg.
12 |  | With 3/4 wheels to move the table and it is easy to use the hammer to lock the wheels anchored to the floor
13 | Accessories:
13.1 |  | 1 X Mattress
13.2 |  | 1 X sliding st/st bowl
13.3 |  | 2 X Leg rest with straps
13.4 |  | 2 X Arm Support
13.5 |  | 2 X Side rails
13.6 |  | 1 X Adjustable IV stand.
13.7 |  | 2 X Shoulder support.
13.8 |  | 1 X Placenta basin made of Stainless Steel.
13.9 |  | 2 X Grab handles.
13.1 |  | 1 X Double Foot Steps
14 | Estimated life-Span of the Item: | At Least 7 Years
15 | Warranty: | Minimum of 2 year.
16 | Other Specification: | Please specify other specification
ITEM No 2 | Cardio Tocogragh CTG, with All Accessories
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier: | the supplier is required to be a certified distributor or agent of the company.
6 | Type: | Mobail on stande with Four casters with brakes.
7 | Control: | The system should be Microprocessor based Foetal Monitor providing continuous monitoring of foetal heart rate (FHR) alongwith maternally sensed foetal activity during antepartum testing for NST
8 |  | Cardiotocograph (CTG monitor), also refereed to as an antepartum foetal monitor). This tabletop unit includes normal and twins Foetal Heart Rate (FHR) and Uterine Contraction (UC) detection, is AC and battery powered and fitted with a loudspeaker, an event marker, a printer and supplied including accessories
9 | Transducer :
9.1 | Type: | Multicrystal wide –beam transducer.
9.2 | Technique : | Autocorrelation.
9.3 | Quantity : | 2 nos (FHR 1, FHR 2)
9.4 | Frequency : | 1MHz to 2 MHz.
9.5 | Intensity : | Less than 10mW/Sq cm.
9.6 | Resolution: | 1 BPM.
9.7 | Heart Rate counting Range: | 30 to 250 BPM.
10 | Printer:
10.1 |  | Facility to print on Inbuilt thermal printer (On thermal paper)
10.2 | Paper : | Z –fold Pre-printed chart scale.
10.3 | Speeds: | 1 to 3cm/ minute.
11 | Features:
11.1 |  | Twin fetal monitoring with TOCO transducers – 01 No.
11.2 |  | It should have clinical event marker.
11.3 |  | It should have monitoring of Bradycardia & Tachycardia alarm events.
11.4 |  | It should have facility to control the volume of FHR sound.
11.5 |  | It should have battery back up of 4-6 hours.
12 | Accessories:
12.1 |  | 1 X Plastic protective dustcover.
12.2 |  | 1 X UC transducer.
12.3 |  | 2 X FHR ultrasound transducers.
12.4 |  | 1 X Remote switch event marker with cable.
12.5 |  | 3 X Adjustable transducer belts for ultrasound (2 FHRs) and toco (UC).
12.6 |  | 2 X Box of thermal recording paper, total 100 z-folded sheets.
12.7 |  | 2 X Bottle of ultrasound gel, approximately 250ml.
12.8 |  | 1 X Mobile cart (Trolley) can be included.
12.9 |  | 1 X Back up battery can be included.
12.1 |  | 1 X Operation Manual
12.11 |  | 1 X Service Manual
13 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
13.1 | Internal back-up battery | It should have battery back up of 2-3 hours.
14 | Essential Requirement:
14.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
14.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
14.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
14.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
14.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
14.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
14.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
14.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
14.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
14.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
14.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
15 | Estimated life-Span of the Device: | At Least 10 Years
16 | Maintenance: | 2 Years free maintenance.
17 | Training: | User /Nurses training, by Specialist from the Supplier.
18 | Other Specification: | Please specify.
ITEM No 3 | Infant Warmer
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier: | the supplier is required to be a certified distributor or agent of the company.
6 | Type: | Mobile on swivel castors , 2 (locks).
7 | Control Unit:
7.1 | Heat Control, Type: | Automatic or manual
7.1.1 | Range, °C: | 34-37 °C Or Better
7.2 | Heater indicator: | Required
8 | Temperature:
8.1 | Set-point Display: | Please specify
8.2 | Skin-temp Display: | Please specify
9 | Alarms:
9.1 | Temperature:
9.1.1 | Tone: | Variable or fixed
9.1.2 | Type: | Audible and visual
9.1.3 | Trigger from setpoint, ±°C/°F | 0.5-1
9.2 | Manual Mode: | Continuous until corrected, automatic shut off if not fixed within 10 min
9.3 | Sensor Disconnect: | Audible and/or visual
9.4 | Power Failure: | Audible and/or visual
10 | Bassinet: | Required
11 | Radiant Heat Source: | Ceramic
12 | Examination Light: | LED
13 | X-ray tray: | Required
14 | Drawers/shelves: | Required
15 | Timer: | Required
16 | Mattress: | Required
17 | Accessories:
17.1 |  | 1 X IV Pole.
17.2 |  | 2 X Spare Heaters
17.3 |  | 2 X Oxygen Holder.
17.4 |  | 3 X Fuses.
17.5 |  | 2 X Spare Led Bulb
17.6 |  | 1 x Operation Manual
17.7 |  | 1 x Service Manual
18 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
19 | Essential Requirement:
19.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
19.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
19.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
19.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
19.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
19.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
19.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
19.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
19.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
19.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
19.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
20 | Estimated life-Span of the Device: | At Least 10 Years
21 | Maintenance: | 2 Years free maintenance.
22 | Training: | User /Nurses training, by Specialist from the Supplier.
23 | Other Specification: | Please specify.
ITEM No 4 | Infant Incubator
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, New Model.
5.2 |  | The Machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the Company.
5 | Type: | Mobile on swivel castors ,2 (locks).
6 | Temperature control range for air, °C: | 23-37°C with (0.1) °C increments and
6.1 |  | Override up to 39 °C
7 | Temperature control range for skin, °C: | 32-37°C with (0.1) °C increments
8 | Temperature Display:
8.1 | Air , display type: | Please specify
8.2 | Range , °C: | 20-42 °C
9 | Temperature Display:
9.1 | Skin , Display Type: | LCD or similar not CRT with a size of min. 5" Or Better
9.2 | Range , °C: | 30-42 °C
10 | Alarms:
10.1 |  | High air temperature
10.2 |  | Low air temperature
10.3 |  | High skin temperature
10.4 |  | Low skin temperature
10.5 |  | Fan failure
10.6 |  | Sensor failure
10.7 |  | Power failure
11 | Front Panel:
11.1 |  | Heater power indicator included
11.2 |  | Controls & switches should be visible, clearly labeled, protected against accidental changes and sealed against penetration of fluids
12 | Design:
12.1 |  | Compact design and heavy duty
12.2 |  | Backup over-heat protection included
12.3 |  | Supplemental humidity adjustable
12.4 |  | Air bacterial filter (changeable)
12.5 | Scale: | Yes
13 | Internal Noise, dB: | ≤ 50 dB approx.
14 | Construction :
14.1 |  | Double wall
14.2 |  | Transparent hood for phototherapy and X-ray
14.3 |  | Movable tray with mattress
14.4 |  | Four antistatic swivel castors, two at least with breaks
14.5 |  | Drawer and instrument tray included
15 | Mattress:
15.1 |  | Adjustable to the trendelenburg and fowler position (Anti trendelenburg)
15.2 |  | Mattress height is adjustable up and down
16 | Ports:
16.1 |  | Hand ports
16.2 |  | Tubing ports:
16.3 |  | O2 inlet ports
17 | O2 Controllers: | Included ( servo Oxygen controllers)
18 | Backup Thermostat: | Required
19 | Heater Power Indicator: | Required
20 | X-Ray Tray: | Please specify
21 | Accessories:
21.1 |  | 1 X Internal Scale Digital Type
21.2 |  | 1 X IV Pole
21.3 |  | 3 X Skin temperature
21.4 |  | 1 x Operation Manual
21.5 |  | 1 x Service Manual
22 | Power Supply: | 100 - 240 V, 50 / 60Hz
23 | Essential Requirement:
23.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
23.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
23.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
23.4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
23.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
23.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
23.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
23.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
23.9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
23.1 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
23.11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
24 | Estimated life-Span of the Device: | At Least 10 Years
25 | Maintenance:
25.1 |  | 2 years free maintenance.
26 | Training: | User /Nurses training, by Specialist from the Supplier.
27 | Other Specification: | Please Specify.
ITEM No 5 | X-Ray Viewer
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The Item Must have lattest Version
5.3 | Quality: | High Quality,
5.4 | Supplier: | the supplier is required to be a certified distributor or agent of the company.
6 | Type: | Wall mounted X-Ray film viewer
7 | Construction: | made from acrylic translucent diffusing materials . screen thickness ≥ 3mm approx. automatic gripping mechanism included
8 | Functionality: | For general radiology applications. Appropriate for viewing a board band rang of film sizes
9 | Viewing panel size (L x W) , cm: | (35 x 43)
10 | Light: | LED
11 | Number of Lamps: | Please specify
12 | On – Off Switch: | Included
13 | Power Supply: | 220 V / 50 Hz
14 | Warranty: | Minimum of 2 years
15 | Other Specification: | Please specify other specification
ITEM No 6 | Ambu Bag / Resuscitation Set (Infant)
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer | Please specify manufacturer and country of origin
2 | Model Number | Please specify model number of the offered equipment
3 | Country Of origin | Please specify Country of origin of the offered equipment
ITEM No 7 | Ambu Bag / Resuscitation Set (Child)
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer: Ambu | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
ITEM No 8 | Ambu Bag / Resuscitation Set (Adult)
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer: Ambu | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
ITEM No 9 | Delivery set Normal + Forceps
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer: Ambu | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Material | stainless steel
5 | component | 1 Pc Bowl Large Size 5″ (12.7cm)
1 Pc Kidney Tray 10″ (25cm)
1 Pc Gallipot 4″ (10.2cm)
1 Pc Sponge Holding Forceps 9.5″ (24.13cm)
1 Pc Scalpel Handle No. 3
1 Pc Needle Holder 8″ (20.3cm)
1 Pc Episiotomy Scissors
1 Pc Umbilical Cord Scissors
2 Pc Spencer Well Artery Forceps Straight 8″ (20.3cm)
2 Pc Kocher Forceps Straight Long 7″ (17.78cm)
2 Pc Spencer Well Artery Forceps Str. 6″ (15.2m)
1 Pc Plain Thumb Dressing Forceps 8″ (20.3cm)
1 Pc Outlet Forceps
ITEM No 10 | SUTURE SET
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer: Ambu | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Material | stainless steel
5 | component | Kidney Dish 20 cm (8") - medium
Gallipot 150 ml - 6 oz - medium
Mayo Hegar needle holder - 18 cm
Kocher artery forceps - straight - 16 cm
Kocher artery forceps - curved - 16 cm
Dressing Forceps (Thumb) - 16 cm
English TOE dissecting forceps - plain - 14 cm
Dissecting forceps toothed 1x2 - 14 cm
Suture scissors - Heaths Ligature - 15 cm
Mayo scissors - straight - 15 cm
Rampley sponge holding forceps - str. - 18 cm
Instrument tray with lid 30 x 20 x 5 cm
ITEM No 11 | DRESSING SET
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer | Please specify manufacturer and country of origin
2 | Model Number | Please specify model number of the offered equipment
3 | Country Of origin | Please specify Country of origin of the offered equipment
4 | Material | stainless steel
5 | component | Kidney Dish 20 cm (8") - medium
Gallipot 150 ml - 6 oz - medium
Kocher artery forceps - straight - 16 cm
Dressing Forceps (Thumb) - 16 cm
English TOE dissecting forceps - plain - 14 c
Mayo scissors - straight - 15 cm
Rampley sponge holding forceps - str. - 18 cm
Instrument tray with lid 25 x 30 x 6.2 cm
ITEM No 12 | D.C SHOCK device
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer | Please specify manufacturer and country of origin
2 | Model Number | Please specify model number of the offered equipment
3 | Country Of origin | Please specify Country of origin of the offered equipment
4 | Technical Specifications | High resolution LCD monitor, ECG recorder and ECG monitor separte ECG input, freeze function ,high resolution printer, synchronous and asynchronous modes,internal rechargable battery,visual and audible alarms Maxoutput 260-360 joule Max charging time to 360joule:5-8 seconds ECGmonitoring time 10 hr Battery charging time to full 4hr With automated external Defibrillition (A.E.D) modewith it's conections& electrode System and malfunction alarms with visual and audible alarm indicators Supplied with Reusable adult pedals ans children pedals & trolly shouls be supplied with all accesories and all connections for adult child and another one for pediatrics Must be certified with CE or FDA markets high quality 2 Years warranty with spear parts from the date of installation and commissioning Technical data set must accombany the offers User/Nurses training by spicialist from the splier
5 | System Configuration | System Configuration Accessories, spares and consumables Defibrillator with AED and External Pacemaker – 01 Adult with Built in Paediatric External Paddles - 01 Patient cables - 01 ECG Rolls – 50 Adult SpO2 reusable Sensor – 01 Adult NIBP Cuff and Hose – 01 etCO2 Tubing (box of 20) – 01 box AED Multifunction Pads for Adults - 10 pairs with Each unit
6 | Environmental factors | The unit shall be capable of operating continuously in ambient temperature of 5 – 45 deg C and relative humidity of up to 95% Shall meet IEC-60601-1-2:2001(Or Equivalent BIS) General Requirements of Safety for Electromagnetic Compatibility.
6 | Power Supply | Power input to be 120-240VAC, 50-60 Hz Should have a battery capable of usage for at least five hours.
ITEM No 13 | HbA1c Analyzer
No. | Technical Specifications | Min MOH Requirement | Suppliers' Conformation Yes / No | If No Clarify
1 | Manufacturer | Please specify manufacturer and country of origin
2 | Model Number | Please specify model number of the offered equipment
3 | Country Of origin | Please specify Country of origin of the offered equipment
4 | Spesification | Color Touch screen portable HbA1c analyzer /eAG POCT fast results Dry system maintenance free Multi-color touch screen Disposable strip, point of care Test paremeters: HbA1c, eAG Sample type : whole blood Sample volume : 3µL-10µL Test time : 3 - 4 min. Measurment range :HbA1c 10%-15.0% or better CV: ≤5 % or better Comparison: 98% Storage :10000 sample; 12 months QC value, power failure protection Power requirment :AC220V±10%, 50Hz,50VA,with battery.
Company Name اسم الشركة |  |  | Signature & Stamp التوقيع و الختم
Representative Nameاسم ممثل الشركة
Designation المسمى الوظيفي
Date: التاريخ