## Technical Specification
Annex B - Technical Specification نموذج المواصفات الفنيه Technical Specifications Tables For The Purchase Medical Equipments for Health Facilities In Al-Hodeidah ,Saadah, and Hajjah Governorates,ADRA-YEMA II-0273-2024.
Technical Specifications
1 |  | Item Description | Electric Blood Bag Tube Sealer
No. | Sub No. | Technical Specifications | Minimum Requirement | Supplier's Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Machine Should be of latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality,Safety, Ease of Use
6 |  |  | Blood Bag Tube Sealer is a Compact Equipment to Seal the Blood Bag Pilot Tubing.
7 |  |  | The System Should be Heavy duty and be Able to Seal the Blood Bag Tubing Quickly and Effectively.
8 |  |  | Should be Simple to Handle.
9 |  |  | System Should Gently Seal the Tubing with no Hemolysis Using Radio Frequency.
10 |  | Wide Sealing: | 2 mm Thickness.Or Better.
11 |  | Sealing Time: | Less Than 5 Sec.
12 |  | Sealing Trigger : | Automatic.
13 |  | Indication Lamps for Sealing Process:. | Indication Lamps
14 |  | No Warm-up Time: | Required.
15 |  |  | Should ensure easy separation of Tube Segments After The Sealing.
16 |  | Back up Battery : | should seal more than 500 seals on PVC tubes in continuous mode.
17 |  |  | Electrodes should be well protected by a cover.
18 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 5 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
19 |  | Life Span for Machine: | 10 Years
20 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Spare Parts list with code NO
21 |  | Training:
| 1 |  | User /Nurses Training, By Specialist From the Supplier.
22 |  | Guarantee:
| 1 |  | The Vendor Should Guarantee The Service and Spare Support for 10 Years of The System and all Accessories after 2 Years of Warranty
23 |  | Power Requirments: | Power Input to be 220-240VAC, 50/60Hz
24 |  | Other Specification: | Please Specify Other Specification.
2 |  | Item Description | Blood Bank Refrigerator
No. | Sub No. | Technical Specifications | Minimum Requirement | Supplier's Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
4.1 | 1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
5.1 | 1 | Designed: | Heavy Duty Designed, New Model.
5.2 | 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
5.3 | 3 | Quality: | High Quality, Safety, Ease of Use
5.4 | 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | General Description: | Blood Bank Refrigerator
7 |  | Maintains Internal Air Temperature: | Between +2°C and +6°C Or Better.
8 |  |  | Blood Bank Refrigerator Is used to store blood bags under controlled temperature.
9 |  |  | Temp Range-should have adjustable termperature control range from +2 °C to +6 °C, factory preset at 4 °C Or Better.
10 |  | Capacity: | Should accommodate 100-200 or More units blood and storage internal volume should be at least 200-300 liters.
11 |  | Refrigerator system:
| 1 |  | The System should have high density CFC -free urethane foam insulation to protect cabinet from ambient temperature fluctuation.
| 2 |  | The system should have positive, forced, air circulation to maintain temperature uniformity at all shelf levels, with quick recovery +/- 1 degree C
| 3 |  | The system should have sensors for activating automatic defrost cycle to minimize the frost build up.
| 4 |  | The system should have automatic condensate removal with no requirement for separate drainage lines.
| 5 |  | Must have fully extendable drawer slides.
12 |  | Construction: | Internal construction should be made up of high grade stainless steel (min 22 G). External construction Corrosion resistant sheet, at least I mm thickness.
13 |  | Internal Temp Control :
| 1 |  | System should have temperature control range from +2 °C to +6 °C.
| 2 |  | Temperature control resolution should be better than 0.1 degree C.
| 3 |  | Cooling down time of max of 120-150 min on half load.
| 4 |  | Should be able to have connectivity to computer and data logger.
| 5 |  | Door System should be lockable with double glass doors for better safety
14 |  | Saftey System:
| 1 |  | System should have large and clear Digital displays for the set/run parameters.
| 2 |  | The system should have weekly chart recorder for temperature changes
| 3 |  | The system should have key operated set point for the added security.
15 |  | Alarm:
| 1 |  | System should have audible/visual warnings for over-temperature, under temperature and power failure with visual status reports on critical functions.
| 2 |  | System should have battery backup and connections for remote alarm contacts.
| 3 |  | Should have adjustments for uneven bases. The adjustments should be easy to use like rotating a screw at the legs In the base.
| 4 |  | Scratch resistant internal lining of the cabinet (stainless steel or aluminium).
16 |  | Display: | Externail tempreture display should be attached, position adjustable
17 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
17 |  | Life Span for Machine: | 10 Years
18 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts list with code NO
19 |  | Accessories:
| 1 |  | 4-5 Removable Shelves.
20 |  | Line Power , VAC: | 100-240, 50/60 Hz
21 |  | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
22 |  | Other Specification: | Please Specify Other Specification.
3 |  |  | Cardio Tocogragh CTG, with All Accessories
No. | Sub No. | Technical Specifications | Minimum Requirement | Supplier's Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 |  | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | high Quality,Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Type: | Mobail on stande with Four Casters with Brakes.
7 |  | Control: | The system should be Microprocessor based Foetal Monitor providing continuous monitoring of foetal heart rate (FHR) alongwith maternally sensed foetal activity during antepartum testing for NST
8 |  |  | Cardiotocograph (CTG monitor), also refereed to as an antepartum foetal monitor). This tabletop unit includes normal and twins Foetal Heart Rate (FHR) and Uterine Contraction (UC) detection, is AC and battery powered and fitted with a loudspeaker, an event marker, a printer and supplied including accessories
9 |  | Transducer :
| 1 | Type: | Multicrystal wide –beam transducer.
| 2 | Technique : | Autocorrelation.
| 3 | Quantity : | 2 nos (FHR 1, FHR 2)
| 4 | Frequency : | 1MHz to 2 MHz.
| 5 | Intensity : | Less than 10mW/Sq cm.
| 6 | Resolution: | 1 BPM.
| 7 | Heart Rate counting Range: | 30 to 250 BPM.
10 |  | Printer:
| 1 |  | Facility to print on Inbuilt thermal printer (On thermal paper)
| 2 | Paper : | Z –fold Pre-printed chart scale.
| 3 | Speeds: | 1 to 3cm/ minute.
11 |  | Features:
| 1 |  | Twin fetal monitoring with TOCO transducers – 01 No.
| 2 |  | It should have clinical event marker.
| 3 |  | It should have monitoring of Bradycardia & Tachycardia alarm events.
| 4 |  | It should have facility to control the volume of FHR sound.
| 5 |  | It should have battery back up of 4-6 hours.
12 |  | Accessories:
| 1 |  | 1 X Plastic protective dustcover.
| 2 |  | 1 X UC transducer.
| 3 |  | 2 X FHR ultrasound transducers.
| 4 |  | 1 X Remote switch event marker with cable.
| 5 |  | 3 X Adjustable transducer belts for ultrasound (2 FHRs) and toco (UC).
| 6 |  | 2 X Box of thermal recording paper, total 100 z-folded sheets.
| 7 |  | 2 X Bottle of ultrasound gel, approximately 250ml.
| 8 |  | 1 X Mobile cart (Trolley) can be included.
| 9 |  | 1 X Back up battery can be included.
| 10 |  | 1 X Operation Manual
| 11 |  | 1 X Service Manual
13 |  | Power Supply: | 100 to 240 V ±10%, 50 Hz.
| 1 | Internal back-up battery | It should have battery back up of 2-3 hours.
14 |  | Essential Requirement:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
| 10 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
| 11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
15 |  | Estimated life-Span of the Device: | At Least 10 Years
16 |  | Maintenance: | 2 Years Free Maintenance.
17 |  | Training: | User /Nurses Training, by Specialist From the Supplier.
18 |  | Other Specification: | Please specify.
4 |  | Item Description | Centrifuge Tabletop
No. | Sub No. | Technical Specifications | Minimum Requirement | Supplier's Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine should be of Latest Model and Must Have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is required to be a Certified Distributor or Agent of The Company.
6 |  | Rotational Speed :
| 1 | Maximum revolutions per minute (RPM) | 4000-5000 Or Better.
| 2 | Maximum relative centrifugal force (RCF), g: | 1100 or More.
7 | 7 | Rotors Available:
| 1 | Fixed Angle: | Included
8 |  | Capacity:
| 1 | Maximum number of tubes and volumes: | 8-12 X (15 mL) Or Better
9 |  | Timer:
| 1 | Range, min: | 1-999 min 59 sec
| 2 | Settings: | Timed, continuous, pulse (burst)
| 3 | Braking Type: | Electric
| 4 | Settings: | 9 or less
| 5 | Brake time, sec | rotor dependent. Please Specify.
| 6 | Brush/Brushless | Brushless
10 |  | Safety Features:
| 1 | Lid Interlock: | Included
| 2 | Alert Indicators:
| 3 | Centrifuge lid Open: | Included
| 4 | End of Run: | Included
| 5 | Overspeed: | Included
| 6 | Display Type: | LCD
| 7 | Parameters Displayed: | RPM, Run Time
| 8 | Lock icon | Included
12 |  | Recommended cleaners/disinfectants: | Ethanol, n-propanol, ethyl hexanol, anionic tensides, corrosion inhibitors
13 |  | Noise Level, Db: | ≤62, Or Less (rotor dependent)
14 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
15 |  | Life Span for Machine: | 10 Years
16 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
17 |  | Power Supply: | 100 to 240 V ~ ±10%, 50 Hz
18 |  | Other Specification: | Please Specify Other Specification.
5 |  | Item Description | Diagnostic Set (ENT) Otoscope & Ophthalmoscope.
No. | Sub No. | Technical Specifications | Minimum Requirement | Supplier's Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates.
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Type: | Portable ENT Diagnostic Set.
7 |  | Main Unit: | Head mirror with adjustable band and 19 mm Diameter.
8 |  | Accessories:
| 1 |  | Ophthalmoscope.
| 2 |  | Otoscope carrier.
| 3 |  | Ear Spacule (3,4,5 & 6 mm diameter.).
| 4 |  | Nasal speculum (8 mm. diameter).
| 5 |  | Dilating nasal speculum.
| 6 |  | Laryngeal mirror stem.
| 7 |  | Laryngeal mirror. Tongue depressor.
| 8 |  | Battery and 2 spare lamps.
| 9 |  | Portable Case.
9 |  | Life Span for Item: | 7 Years
10 |  | Warranty: | Minimum of 2 Years.
11 |  | Other Specification: | Please Specify Other Specification.
6 |  | Item Description | Fetoscope - Audio Fetal Doppler
No. | Sub No. | Technical Specifications | Minimum Requirement | Supplier's Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | 5 | Genaral Specification: | The Unit shall be safe to use
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Type: | handheld.able for high load & hard work
7 |  | General Description:
| 1 |  | Foetal Heart Rate (FHR) doppler detector, with speaker, handheld, battery powered, with accessories
8 |  | Sensitivity From : | 10 - 12 weeks Or Better
9 |  | Technical Specifications:
| 1 |  | Foetal Heart Rate (FHR) ultrasound detector, single transducer
| 2 |  | Handheld lightweight detector
| 3 |  | Ultrasonic continuous doppler operating mode
| 4 |  | External transducer, connected to the main unit through a coiled wire cable
| 5 |  | Integrated transducer holed on the body of the main unit
10 |  | Ultrasound frequency : | 2 MHz ± 10 %
11 |  | Sensitivity : | 90 dB
12 |  | Ultrasonic Power Output : | mW, depending on the available model
13 |  | FHR Range : | 50 – 240 bpm ± 2 bpm, resolution 1 bpm
14 |  |  | Integrated loudspeaker and volume level adjustment
15 |  |  | Noise suppression feature
16 |  |  | Integrated headset connector
17 |  | Display: | LCD display, backlighted, with FHR, pulse, frequency, volume, battery and system status
18 |  | Alarms: | Audio/visual alarms for high/low FHR, battery status and system errors
19 |  |  | Automatic Self-Test
20 |  |  | Automatic Power-Off
21 |  | Accessories:
| 1 |  | 2 X 350 ml or 3 x 250 ml (depending on availability) bottles of US gel
| 2 |  | 2 X sets of commercial disposable alkaline batteries (separately packed)
| 3 |  | 1 X soft carry bag
| 4 |  | 1 X Instructions for assembly, use and maintenance in English
22 |  | Battery Autonomy: | >5 h or 600 exams, whichever higher
23 |  | Power Supply: : | 1.5 V AA (LR6) Commercial disposable batteries or 1.2 V Ni-MH commercial rechargeable batteries (not included), 2 or 3 units depending on the available model
24 |  | Essential Requirement:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
| 10 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
| 11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
25 |  | Estimated life-Span of the Device: | At Least 5 Years
26 |  | Maintenance: | 2 Years Free Maintenance.
27 |  | Training: | User /Nurses training, by Specialist from The Supplier.
28 |  | Other Specification: | Please Specify Other Specification.
7 |  | Item Description | Defibrillator With ECG Monitor
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | 5 | Genaral Specification: | The Unit Shall be Safe to use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Defibrillator Technology : | Biphasic
7 |  | Defibrillator :
| 1 | Output Energy Range: | 5-260J Or Better , to be selected through not less than 10 steps
| 2 | Internal Discharge: | Included
| 3 | Charging Trigger: | On paddles
| 4 | Synchronizer: | Included
| 5 | Max. Energy Charge Time: | Max. 10 sec.
| 6 | Operation Modes: | Sync. , Async.
| 7 | Discharge Test: | Included
| 8 | Charging Indication: | Audio and visual
| 9 | AED: | With Automated External Defibrillation ( A.E.D ) Mode with it's connections & Electrode .
8 |  | ECG Monitor :
| 1 | Display: | LCD or similar not CRT with a size of min. 5" Or Better
| 2 | Displayed Information: | Heart rate , ECG waveform , ECG source , output energy selected and delivered
| 3 | ECG Source: | Manually selected , paddles , leads
| 3.1 |  | I , II , III
| 4 | 1 mV Cal. Signal: | Included
| 5 | Heart Rate Range: | 30-250bpm with audio and visual alarm with selected High and Low limits
| 6 | ECG Gain: | 5 , 10 , 20 mm/mV
| 7 | Heart Beat Indicator: | Audio and visual
9 |  | Recorder:
| 1 | Type: | Built in thermal head printer
| 2 | Speed: | Adjustment 5 - 25 mm/sec.
| 3 | Modes: | Auto & Manual
10 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
11 |  | Life Span for Machine: | 10 Years
12 | 1 | Maintenance:
| 2 |  | 2 Years Free Maintenace.
| 3 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 4 |  | Application Software and Interface Connection Included.
| 5 |  | Spare Parts list with code NO
13 |  | Accessories :
| 1 | Trolley | High quality trolley on 4 castors , 2 with brakes with drawer for the accessories
| 2 | Paddles: | External for Adult & Pediatric
| 2.1 | ECG Cable: | 1 X 3/5 Lead wire cable for adult
| 2.2 |  | 1 X 3/5 Lead wire cable for pediatric
| 3 | All other needed accessories | Included ( 3 Roll Paper)
14 |  | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
15 |  | Rechargeable Battery: | Rechargeable battery with built in charger with indicators for charging and low battery
16 |  | Battery: | At Least 2 Hrs.
17 |  | Other Specification: | Please Specify Other Specification.
8 |  | Item Description | Digital Glucometer
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
| 2 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  |  | Digital Glucometer.
7 |  | Technical Specifications:
| 1 |  | Sample volume one drop
| 2 |  | Test Time Less than 10 seconds
| 3 |  | With Memory
| 4 |  | Set complete with prick pen
| 5 |  | Lancets
| 6 | Measurement Unit : | mmol/L Or mg/dL
| 7 | Sample Volume : | 0.5 microliters
| 8 | Display : | LCD Display
| 9 |  | Battery Operated
8 |  | Accessories:
| 1 |  | 500 X Glucose Strips
| 2 |  | 1 X Carry Bag.
| 3 |  | 2 X Battery (Spare part)
9 |  | Life Span for Machine: | 10 Years
10 |  | Warranty: | Minimum of 1 Year.
11 |  | Other Specification: | Please Specify Other Specifications.
9 |  | Item Description | Gynecological Delivery Bed ( Hydraulic Type)
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Type: | Hydraulic Type
7 |  |  | For Delivery & Obstetric Use
8 |  | Materials: | Three sections,Made of steel, Epoxy-Painted with sturdy sampling of material .or any Equivalent rust-proof Material
| 1 |  | Heavy duty ABS mattress base, 80 mm durable hygienic Mattress,
| 2 |  | collapsible side rails, leg section mattress, adjustable foot supports, retractable foot section
9 |  | Adjustments:
| 1 | Height: | from 550 to 900 mm
| 2 | Back section: | 70 degree gas spring assisted
| 3 | Trendlenburg: | (+ 12 degree & -3 degree)
10 |  |  | With the adition of Pelvic position
11 |  | Permissible Load: | 120 - 150 Kg.
12 |  |  | With 3/4 wheels to move the table and it is easy to use the hammer to lock the wheels anchored to the floor
13 |  | Accessories:
| 1 |  | 1 X Mattress
| 2 |  | 1 X sliding st/st bowl
| 3 |  | 2 X Leg rest with straps
| 4 |  | 2 X Arm Support
| 5 |  | 1 X Adjustable IV stand.
| 6 |  | 2 X Shoulder support.
| 7 |  | 1 X Placenta basin made of Stainless Steel.
| 8 |  | 2 X Grab handles.
| 9 |  | 1 X Double Foot Steps
14 |  | Estimated life-Span of the Item: | At Least 7 Years
15 |  | Warranty: | Minimum of 2 Years.
16 |  | Other Specification: | Please Specify Other Specification
10 |  | Item Description | Patient Bed Three Sections With Locker, Bedside and Over Bed Table
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Design & Quality: | Bed, designed 3 Sections standard wards .
7 |  | Construction: | Epoxy-Coated main frame constructed from heavy duty section steel tube
| 1 |  | Removable Chrome-Plated steel bed ends fitted with 10mm laminated plastic panels
| 2 |  | 3 Sections mattress Platform fitted with Epoxy-Coated steel weldmesh
8 |  |  | Backrest and knee brake adjustable by hydraulic action
9 |  |  | The bed shall be mounted on 4x125mm (minimum) castors, 2 swivelling with simultaneous brakes
10 |  |  | movable bed sids,four five- inch castors two of them are totally lockable,adjustable
11 |  | Guiding Dimentions of bed: | 210-220L x 90-100Wx 30-60H cm
| 1 | Accessories:
| 2 |  | Three sections foam. Mattress, 10-12cm at least High density, Premium-grade cotton felt and high-density urethane foam are intertufted to eliminate component shifting and to increase mattress life with Durable vinyl cover, anti-bacterial, anti-static, acid-resistant and waterproof for easy use and care
| 3 |  | IV rod, 2 hooks with adjustable height
| 4 |  | Side folding-down rails (pair)
| 5 |  | 1 X Pillow With Cases
| 6 |  | 1 X Hospital Sheet
12 |  | Locker, Bedside | Heavy duty, Flexible, high quality
| 1 | Construction: | Made from the same Material and same Colors of bed,better From the same Beds Manufacturer
| 2 |  | 4 castors (Orientable)
| 3 |  | Upper drawer, with nylon ball bearing sliders. Internal Plastic receiver
13 |  |  | Over Bed Table Heavy duty, Flexible &Foldable, high quality
| 1 | Over Bed Table Intgreted or Sparately: | Please Specify
| 2 |  | Gas Spring Height Adjustment By Lever.
| 3 |  | Table Top Made Of Playwood.or equivalen Material for easy Cleaning
| 4 | Height Adjustment Range : | 700-1200 mm.
15 |  | Life Span for Item: | 7 Years
16 |  | Warranty: | Minimum of 2 Years For All.
17 | 17 | Other Specification: | Please Specify Other Specification.
11 |  | Item Description | Infusion Pump W/Holder
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 |  | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Mode:
| 1 |  | Macro : 1-999 ml/hr in 1ml./hr increment Or Better.
| 2 |  | Micro : 0.1-99.9 ml/hr in 0.1 ml/hr increment Or Better.
7 | 7 | Volume to be infused: | 1.0 – 999.9 ml or No limit Or Better.
8 |  | Infusion Set Used with the System: | should (open system) the user can use any infusion set
9 |  | Infusion Accuracy:
| 1 |  | Presettable rate
| 2 |  | At Least ±2% Or Better.
| 3 | Secondary Function:
10 | 1 |  | Free flow protection
| 2 |  | Front panel key pad lock
| 3 |  | Built-in self test and diagnostic
| 4 |  | System and malfunction alarms with audible and visual alarm indicators
| 5 |  | Silence or reset of audio alarms
| 6 |  | Air line detect system
11 |  | Message Display for Infusion Status: | Alarms, error messages etc.
12 |  | Automatic Flow:
| 1 |  | Rate calculation
| 2 |  | Drug calculation & drug dose programs
| 3 |  | Multi-dose capability
| 4 |  | Rate-volume and volume timer programming
13 |  | Accessories:
| 1 |  | 1 X Holder
| 2 |  | 1 X Operation Manual
| 3 |  | 1 X Service Manual
| 4 |  | 1 X Set of Spare Infusion Set For Adult & Pediatric. (each) 500 Set
14 |  | Power S:upply: | 100 to 240 V ~ ±10%, 50/60 Hz .
| 1 | Battery Pickup: | 4-6 Hrs.
15 |  | Essential Requirement:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
| 10 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
| 11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
16 |  | Estimated life-Span of the Device: | At Least 10 Years
17 |  | Maintenance: | 2 Years free maintenance.
18 |  | Training: | User /Nurses training, by Specialist from the Supplier.
19 |  | Other Specification: | Please Specify Other Specification
12 |  | Item Description | Laboratory Refrigerator
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of the Company.
6 |  | General Description: | Laboratory Refrigerator
7 |  | Maintains internal air Temperature: | Between +2°C and +8°C
8 |  | Construction: | Smooth, molded steel cabinet with rust-proof enamelled finish and bevelled edges
9 |  |  | Easy non-electric push-button lighting. Shall have no corners to trap dirt and bacteria.
10 |  |  | The absorption-type cooling system shall be quiet and dependable, with no compressor or other moving parts to wear out
11 |  | Alarm: | Audio/visual alarm signal on temperature deviation, with remote alarm terminals providing mains failure alarm signal
12 |  |  | Shall be easy-to-clean and all-white interior features adjustable, removable shelves and door bins for stocking flexibility
13 |  | Internal Capacity: | shall be minimum 200 -300 liters.
14 |  |  | Lockable door, 2 keys for security
15 |  | Display: | Digtal Display.
16 |  | Essential Requirement:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
17 |  | Life Span for Machine: | At Least 10 Years
18 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts list with code NO
19 |  | Accessories:
| 1 |  | 4-5 removable shelves
20 |  | Line Power , VAC: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase ( power cable Compatible with the Hospital electric outlet plug, 5 mt ), Electrical Safety class 1
21 |  | Other Specification: | Please Specify Other Specification.
13 |  | Item Description | Laryngoscopes Adult - Pediatric
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Material of Construction: | Heavy duty Stainless Steel 304 or better.
7 |  | Light Type: | LED
8 |  | Accessories:
| 1 |  | 5 X LED lampe Spare 3V
| 2 |  | 4 Set X Different Size for Adult
| 3 |  | 3 Set X Different Size for Pediatric
| 4 |  | 1 X Box Battery
| 5 |  | 1 X Career case for save the instrument
9 |  | Life Span for Item: | 7 Years
10 |  | Warranty: | Minimum of 1 Year.
11 |  | Other Specification: | Please Specify Other Specification.
14 |  | Item Description | Ultrasonic Nebulizer Portable
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The machine shuld be of latest model and must have lattest technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of the Company.
6 |  | Nebulizing Rate, mL/min: | 3 mI / min At Least.
| 1 | Medical fluid Chamber Capacity: | 250-300 ml. External Cup
| 2 | Medication Cup Capacity: | 100-150 ml Or Better, Internal Cup
| 3 | Nebulizing Timer: | 0 - 30 Minutes Or Better.
| 4 |  | Continuous Mode
| 5 | Crystal Operating Frequency : | 1.5-2 Mhz +/- 5%.
7 | 7 | Alarms :
| 1 |  | Buzzer to alert of any errors or when time is completed
8 |  | Type: | Portable
9 |  | Operating pressure | 0-1.0 bar.
10 |  | Noise level: | Please specify.
11 |  | Max. Pressure: | Please specify.
12 |  | Spare Parts: | Recommended spare parts list with price and validity must be provided along with the offer (Mandatory).
13 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
14 |  | Life Span for Machine: | 10 Years
15 |  | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
16 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts list with code No
17 |  | All standard accessories | All accessories (filters, tubings, face mask (Adult,Pediatric,Infant, etc.)
18 |  | Power supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
19 |  | Battery: | At Least 2 Hrs
20 |  | Other Specification: | Please Specify Other Specification
15 |  | Item Description | Syringe & Needle Destroyer
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark | Required Certificates
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Body Material: | ABS Plastic.
7 |  | Type fo Electrodes Used: | Copper Electrode- Nikle plated.
8 |  | Destruction Rate, seconds/Needle : | 5-10 Seconds/ Needle
9 |  |  | Thickness & Length of needle Recommended: Max, 1mm dia, 80mmlength.
10 |  | Grade of Steel & Guaranteed Life of Cutting Blade, (Approx. No. of Syringeand Needle it Can Destroyer): | Steel grade is SS420. Can cut syringe hub &burn needle approx. 25000 pcs. To 30000 pcs. Minimum.
11 |  | Tray Capacity, No. of Syringes & Needles: | 200-300pcs.
12 |  | Working Noise: | ≤50 dB
13 |  | Life Span for Item: | 7 Years
14 |  | Warranty: | Minimum of 1 Year.
15 |  | Other Specification: | Please Specify Other Specification.
16 |  | Item Description | Oxygen Cylinder 40L With Oxygen Regulator, Humidifier and Trolley
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark | Required Certificates
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Type: | Oxygen cylinder Empty.40L (60m3)
7 |  | Medical Oxygen Cylinder:
| 1 |  | Oxygen cylinder spindle key for bull nose, Spare spindle key for pin index
| 2 | Pressure: | Working pressure 150Bar/200Bar
| 3 | Capasity: | 40 liters (Bull Nose British standard).
| 4 | Material: | Special Steel (Anti Rust). Blue Or Black Color.
8 |  | Oxygen Trolley on castor:
| 1 | Quality : | High quality.
| 2 | Material: | Special Steel.
| 3 | Load: | (40-50 kgs cylinder).
9 |  | Oxygen Regulator and Humidifier:
| 1 | Flow Range: | From 0.15 L/min to 15 L/min.
| 2 | Maximum possible flow: | 35 L/mn above range.
| 3 | Input pressure: | Not less than 3.5 bar +/- 0,7
| 4 | Suitable for all types of fitting: | Wall outlet, trolley.
| 5 |  | Autoclavable, Unbreakable
| 6 |  | Clear vision of bubbles
| 7 | Humidification at 4 L/mn - 20° C: | ~3 ml/h
| 8 | Humidification at 12 L/mn - 20° C: | ~9 ml/h
| 9 | Capacity: | 100 ml (distilled water)
| 10 | Porosity of diffuser: | No More than 70 Microns
| 11 |  | Humidifier, Autoclavable, for multipurpose uses in Medical and Surgical Sections. Shall have the capability to be associated with Flowmeters (O2) or Air,O2 Reserve Cylinders and O2 Generators
| 12 |  | Fixation to flowmeters or any compatible devices shall be easy and fast (1/4 twist shall be acceptable)
10 |  | Oxygen masks (for child.&adult) & connections :
| 1 | Quality : | High Quality.
| 2 | Material: | Silicone and Rubber.
11 |  | Accessories:
| 1 |  | 3 X Face Mask diffrent Size for adult. Resuable Type
| 2 |  | 3 X Face Mask diffrent Size for children. Type
| 3 |  | 1 X Tubing.
| 4 |  | 1 x Oxygen Trolley
12 |  | Estimated life-Span of the Item: | At Least 10 Years
13 |  | Warranty: | Minimum of 1 Year.
14 |  | Other Specification: | Please Specify Other Specification.
17 |  | Item Description | Patient Monitor With Holder
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | 5 | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: Class I, Waterproofnessl.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Display size and Type: | 12-15 Inches Or Better. Touch Screen
High resolution multi-color display
Medical type (preferable) or supported with isolation transformer
7 |  | Displayed information: | Min. 6 vital waveforms, cab be colored separately
Numeric data for the measured vital parameters
Vital parameters (72) hrs trends
Vital parameters alarms audio visual
8 |  | Defibrillation Protection: | Available for the measured vital parameters
9 |  | Required vital parameters:
| 9.1 | ECG:
| 9.1-1 | Leads: | 12 leads facility
| 9.1-2 | Values: | ECG lead waveform, label, HR gain as min.
| 9.1-3 | HR Range: | 30-200 bpm
| 9.1-4 | Alarms: | Leads off, Hi + Low HR
| 9.1-5 | Gain: | 5, 10, 20 mm/mV
| 9.1-6 | Preferable items: | ST, Arrhythmia, cascade ECG
| 9.2 | Respiration:
| 9.2-1 | Technology: | Impedance
| 9.2-2 | Values: | Respiration waveform, R.R.
| 9.2-3 | R.R. Range: | 5-80 bpm
| 9.2-4 | Alarms: | Hi & low R.R.
| 9.2-5 | Apnea Alarm: | 15-25 sec. Preferable
| 9.3 | NIBP:
| 9.3-1 | Technology: | Oscillometric
| 9.3-2 | Values: | SYS , DIA , MEAN , P.R.
| 9.3-3 | Modes: | Auto, Manual
| 9.3-4 | Cuff Pressure Range:
| 9.3-4-1 |  | Up to 250mmHg Adult
| 9.3-4-2 |  | Up to 200mmHg Pediatric
| 9.3-4-3 |  | Up to 150mmHg Neonate
| 9.3-4-4 | Alarms: | Hi and low, SYS, DIA
| 9.4 | SPO2:
| 9.4-1 | Values: | SPO2 waveform, SPO2% , P.R.
| 9.4-2 | SPO2 range | 50 - 100 %
| 9.4-3 | Pulse Rate Range: | 25-200ppm
| 9.4-4 | Alarms: | Hi & low SPO2 + P.R. sensor off
| 9.5 | Temperature:
| 9.5-1 | Values: | T1 and or T2
| 9.5-2 | Range: | 30-45°C
| 9.5-3 | Alarm: | Hi & low
| 9.6 | ETCO2:
| 9.6-1 |  | Should be Main Stream Capnography with Display of CO2 and Digital Values of EtCO2, FiCO2 & RR.
| 9.6-2 | EtCO2 Range : | 0-99 mmHg.
| 9.6-3 | FiCO2 : | 0 to 20 mmHg.
| 9.6-4 | Flow rate : | 50ml/min.
10 | 10 | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
11 |  | Life Span for Machine: | 10 Years
12 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts list with code NO
13 |  | Required Accessories:
| 1 |  | 1 X 3/5 Leads wire ECG cable Adult.
| 2 |  | 1 X 3/5 Leads wire ECG cable Pediatric.
| 3 |  | 1 X NIBP reusable cuff large Adult
| 4 |  | 1 X NIBP reusable cuff pediatric
| 5 |  | 1 X NIBP reusable cuff Infant
| 6 |  | 1 X NIBP connection hoses
| 7 |  | 1 X SPO2 reusable finger probe + Extension cable adult.
| 8 |  | 1 X Temperature reusable sensor skin Type for Pediatric and Infant
| 9 |  | 1 X EtCO2 module with all accessories. ( Cable-Sensor)
| 10 |  | 1 X Original Holder In The Wall
| 11 |  | Operation Manual
| 12 |  | Service Manual
| 13 |  | Spare parts list with code NO.
14 |  | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
15 |  | Power Supply: | 100 to 240 V ±10%, 50 Hz.
| 1 | Internal back-up battery | Operating time, hr : ≥2
16 |  | Other Specification: | Please Specify Other Specification.
18 |  | Item Description | Semi Auto Coagulation Analyzer
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | 5 | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Method Used: | Please Specify
7 |  | Features:
| 1 | Principle: | Clotting, Optical method.
| 2 | Light Source : | Lamp Or Led
| 3 |  | Precise electronic pipette
| 4 |  | Automatic start reagent addition
| 5 |  | Thermal Printer, auto & manual print
8 |  | Testing Channel : | (2-4)
9 |  | Display: | At least 5 inch LCD
10 |  | Cuvette Incubation Positions : | 12-24. Or Better
11 |  | Reagent Incubation Positions : | 4-6. Or Better
12 |  | Wavelength : | 470nm.
13 |  | Sample Type : | Plasma.
14 |  | Memory : | At least 10000 results Storage.
15 |  | Sample Volume: | 20-40μl.
16 |  | Electronic Pipette : | 20-40μl
17 |  | Available dimensions : | Seconds, ratio, %, g/L
18 |  | Incubation Temperature °C: | 37.0°C ±1.0℃
19 |  | Deviation between channels : | ≤5%
20 |  | Printer : | Inner Thermel Printer.
21 |  | Reagent Type: | (Open system )
22 |  | Test Menu:
| 1 |  | PT
| 2 |  | APPT
| 3 |  | TT
| 4 |  | Fib
|  | Other Test Menu: | Please Specify
23 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
24 |  | Life Span for Machine: | 10 Years
25 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts list with code NO
26 |  | Accessories:
| 1 |  | One Spare Led Lamp
| 2 |  | Provide free Kit for Operation and Training on the use of the Machine.
| 3 |  | UPS Online 1 KVA
27 |  | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
28 |  | Power Supply: | 100 to 240 V ±10%, 50 Hz.
29 |  | Other Specification: | Please Specify Other Specification.
19 |  | Item Description | Sphygmomanometer - Mercury 300 mm With Cuffs
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 |  | Genaral Specification: | The Unit shall be safe to use
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Type: | Portable Mercurial Type
7 |  |  | Should have ON and OFF provision for mercury reservoir ON and OFF provision for mercury reservoir
8 |  |  | A measuring range from 0 to 300 mmHg.
9 |  |  | The control valve knurled thumb control device. The leak rate should not exceed 10 mm of mercury per minute.
10 |  |  | The manometer scale markings and graduations should be permanent and clearly visible and filled with pigments.
11 |  |  | The internal diameter of the manometer glass tube should be 4.1 ± 0.1 mm and the thickness not less than 2 mm.
12 |  |  | All plastic parts, if any used should not crack, flake, peel or disintegrate in normal use.
13 |  |  | The inflating rubber bag should be capable of withstanding an internal pressure of 450 mmHg without leaking.
14 |  |  | The inflating bulb should be soft and should not have any joints or ridges.
15 |  |  | The mercury used should be clean, double distilled and of 99.9% purity.
16 |  |  | The fastening arrangements of the cuff should be of hook and loop type (Velcro).
17 |  |  | The threading and fastening arrangement of the cuff should show no sign of slip or failure when subjected to the maximum test conditions.
18 |  |  | The rubber tubes used should have an internal diameter of 3 ± 0.5 mm and the external diameter should not be less than 8mm.
19 |  |  | The tubes should be fitted with male and female leur connectors.
20 |  |  | The housing case should be of robust design. It should have press to release lock. It should have metal hinges. The tube should be secured with metal screws and clamps. It should have mechanism to hold the lid in right angles and should prevent accidental dropping. All parts should be replaceable in case of breakage
21 |  |  | A cleaning brush to clean the manometer tube and a set of spare washers may be provided with each unit.
22 |  | Accessories:
| 1 |  | 3 X Adult Arm Cuffs of Size Medium & Large and Small cuff.
| 2 |  | 1 X Child Cuffs Resauble Type
| 3 |  | 1 X Neonate Cuffs Resauble Type
| 4 |  | 1 X Spare Silicon Connecting Tube
| 5 |  | 1 X Storge Case
| 6 |  | 1 X Operation Manual
23 |  | Estimated life-Span of the Item: | At Least 10 Years
24 |  | Warranty: | Minimum of 2 Years.
25 |  | Other Specification: | Please Specify Other Specification
20 |  | Item Description | Bench Top Steam Sterilizer With Drayer
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier | the Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Air Removal Method: | To Facility Requirements
7 |  | Chamber Material: | 314 Stainless Steel Or Better.
| 1 | Type: | Class B
| 2 | Number of trays: | 2-3 Or Better.
6 |  | Water Reservoir Capacity , L: | 40 L
9 |  | Preprogrammed Cycles, Type: | Wrapped, Unwrapped
10 |  | Cycle Number ,Type: | Please Specify
11 |  | Minimum Cycle Time ,Min: | Please Specify
12 |  | Automatic Cycle Shut off: | Included.
13 |  | Temperature Range , °C : | 100-134 °C Or Better.
| 1 | Overheat shutoff: | Included.
14 |  | Pressure Purge: | Please Specify
15 |  | Controller Type: | Please Specify
|  | Programmable: | Included.
16 |  | Alarm Conditions:
| 1 | Indicator type: | Audible, visual
| 2 | Chamber pressure or temperature Failure: | Audible, visual
| 3 | Component Failure: | Audible, visual
| 4 | Cycle disruption or Failure: | Audible, visual
17 |  | Data Management:
| 1 | Recorder: | Included.
| 2 | Printer: | Included.
18 |  | Power Supply, VAC: | 110-230 50/60 Hz
| 1 | Consumption, Kw: | Please Specify
19 |  | Drayer Function: | Included. Vaccum Type
20 |  | Accessories:
| 1 | Steel Container: | plain basket with handles.
| 2 | Paper for Printer: | 10 rolls
| 3 |  | Spare Part Heater 1 Pcs.
| 4 |  | Rubber Gasket for Door 1 Pcs.
21 | 21 | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
22 |  | Life Span for Machine: | 10 Years
23 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
24 |  | Power Supply: | AC 110-240V, 50/60Hz single-phase electrical source.
| 1 |  | Compliant with IEC 60601.
| 2 |  | Built-in protections against over-voltage and over-current line conditions.
25 |  | Other Specification: | Please Specify Other Specification
21 |  | Item Description | Mobile Electric Suction Machine 2X2 L
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The machine shuld be of latest model and must have lattest technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is required to be a Certified Distributor or Agent of the Company.
6 |  | Flow rate L/ Min: | Not less than approx. 30 liter / min.
7 |  | Vacuum , mmHg: | Not less than 700 mmHg.
8 |  | Vacuum regulate: | Manual key adjust the vacuum volume to high or low.
9 |  | Bottles: | Two collecting jars (2) liters fitted with sterilizable rubber bung & fluid valve. Change-over block from jar to other
10 |  | Tube: | Suction tube between operation field and suction machine) reusable length 3m Material silicone, Autoclavable
11 |  | Castors: | The apparatus should be mounted on four anti-static tired rubber castors, Two With Break
12 |  | Filters: | The apparatus should be available with bacterial filters
13 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
14 |  | Life Span for Machine: | 7 Years
15 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
16 |  | Accessories: | 2 Spare bottles, 10 spare bacterial filters and 1 tube (Must be included in the price)
17 |  | Power Supply: | 100 to 240 V ~ ±10%, 50 Hz
18 |  | Other Specification: | Please Specify Other Specifications.
22 |  | Item Description | Syring Pump W/Holder
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK
5 |  | Genaral Specification: | The Unit Shall be Safe to Use Both for -The Operator and The Patient. Class I, Protection Against Leakage Current, Protection Against Electric Shocks: class I, Waterproofnessl.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Syring Pump: | Permits the use of standard disposable Syringes (sizes 10 – 50/60 ml) and extension sets, with auto syringe size sensing.
7 |  | Infusion Volume: | 0.1 to 999 ml in 0.1ml increments in lower ranges
8 |  | Infusion rate adjustable: | from 0.1 – 99.9 ml/hr in (0.1ml/hr steps)
9 |  | System and malfunction: | Alarms with visual and audible alarm indicators
10 |  | Display:
| 1 |  | Digital display or better of volume infused. Infusion rate, alarm prompts etc.
| 2 |  | Purge/ bolus facility
| 3 |  | Volumetric accuracy At Least ± 2%.
| 4 |  | Occlusion Pressure limit (from 0 to 750 mmHg)
| 5 |  | Built-in self test and diagnostics
11 |  |  | DPS (dynamic pressure sensor) Included
12 |  |  | Special feature of the product (other than mentioned on list)
13 |  |  | Accessories if any (to put machine/ system into operation)
14 |  |  | Consumable if any (to put machine/ system into operation)
15 |  | Accessories:
| 1 |  | 1 X Holder
| 2 |  | 1 X Operation Manual
| 3 |  | 1 X Service Manual
16 |  | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
| 1 | Battery: | 4-6 Hrs.
17 |  | Essential Requirement:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | To submit a spare part pricelist valid for 7 years starting date after the warranty completion
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation.
| 10 |  | Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number.
| 11 |  | Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion.
18 |  | Estimated life-Span of the Device: | At Least 10 Years
19 |  | Maintenance: | 2 Years Free Maintenance.
20 |  | Training: | User /Nurses training, by Specialist from the Supplier.
21 |  | Other Specification: | Please Specify Other Specification
23 |  | Item Description | Oxygen Regulator and Humidifier
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark | Required Certificates
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Oxygen Regulator and Humidifier:
| 1 | Flow Range: | From 0.15 L/min to 15 L/min.
| 2 | Maximum possible flow: | 35 L/mn above range.
| 3 | Input pressure: | Not less than 3.5 bar +/- 0,7
| 4 | Suitable for all types of fitting: | Wall outlet, trolley.
| 5 |  | Autoclavable, Unbreakable
| 6 |  | Clear Vision of Bubbles.
| 7 | Humidification at 4 L/mn - 20° C: | ~3 ml/h
| 8 | Humidification at 12 L/mn - 20° C: | ~9 ml/h
| 9 | Capacity: | 100 ml (Distilled Water)
| 10 | Porosity of diffuser: | No More Than 70 Microns
| 11 |  | Humidifier, Autoclavable, for multipurpose uses in Medical and Surgical Sections. Shall have the capability to be associated with Flowmeters (O2) or Air,O2 Reserve Cylinders and O2 Generators
| 12 |  | Fixation to flowmeters or any compatible devices shall be easy and fast (1/4 twist shall be acceptable)
7 |  | Estimated life-Span of the Item: | At Least 10 Years
8 |  | Warranty: | Minimum of 1 Year.
9 |  | Other Specification: | Please Specify Other Specification.
24 |  | Item Description | Laboratory Microscope Binocular
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
1 |  | FDA Certificate Or CE Mark: | Required Certificates
2 |  | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Machine should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | the Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Type: | Binocular
7 |  | Eyepieces: | Plese Specify
8 |  | Interpupillary distance adjustment, mm: | Adjustable
9 |  | Nosepiece Configuration: | Quadruple or better
10 |  | Objectives magnification: | 4x,10x; 40x; 100x (oil)
11 |  | Objectives Type: | Achromatic OR Plan achromatic
12 |  | Filter:
| 1 | Color(s): | Please Specify
13 |  | Illumination:
| 1 | Condenser Type: | Built-in
| 2 | Numerical Aperture: | Please Specify
| 3 | Light Source: | Halogen or Led
14 |  | Stand:
| 1 | Focusing Mechanism: | Moving Stage
| 2 | Coarse, Fine Adjustments: | Coaxial Coarse and Fine Adjustment
15 |  | Stage:
| 1 | Type: | Mechanical
| 2 | Tension adjustment: | Coaxial
| 3 | Motion: | X-Y
16 |  | Accessories:
| 1 | Adapter(s): | Required
| 2 | Dust cover: | Required
| 3 |  | 50 X Microscope Slides 76x26mm.
| 4 |  | 100 X Cover slides 22x22mm.
17 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
18 |  | Life Span for Machine: | 10 Years
19 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
20 |  | Power Supply: | 100 to 240 V ~ ±10%, 50 Hz
21 |  | Other Specification: | Please Specify Other Specifications.
25 |  | Item Description | Water Bath
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Machine should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | General Description: | Water Bath for Laboratory
7 |  | Constructions: | Stainless Steel,
8 |  | Display: | Digital Display.
9 |  |  | Circulating
10 |  | Control: | Electronic control.
11 |  |  | Over temp cut off
12 |  |  | Perforated shelf with cover
13 |  | Capacity: | 7-15 Liters, temperature 25°C-85°C.approx.
14 |  | Temperature : | Ambient+ 5 to 85°C.approx.
15 |  | Timer: | About 30 Minutes Or Better.
16 |  | Sensitivity: | ± 0.2o C or Less
17 |  | Uniformity: | ± 0.1o C or Less
18 |  | Heater power: | 1500 W or Less
19 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
20 |  | Life Span for Machine: | 10 Years
21 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
22 |  | Accessories:
| 1 |  | Complete with Ccover.
| 2 |  | Heater Spare Part
| 3 |  | Holder of Tubes: for 10 tubes
| 4 |  | 3 Fuses
23 |  | Power Supplay: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase ( power cable Compatible with the Hospital electric outlet plug, 5 mt ), Electrical Safety class 1
24 |  | Other Specification: | Please Specify Other Specification.
26 |  | Item Description | Oxygen Cylinder 10L With Oxygen Regulator, Humidifier and Trolley
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark | Required Certificates
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Type: | Oxygen cylinder Empty.10L (15 m3)
7 |  | Medical Oxygen Cylinder:
| 1 |  | Oxygen cylinder spindle key for bull nose, Spare spindle key for pin index
| 2 | Pressure: | Working pressure 150Bar/200Bar
| 3 | Capasity: | 10 liters (Bull Nose British standard).
| 4 | Material: | Special Steel (Anti Rust). Blue Or Black Color.
8 |  | Oxygen Trolley on castor:
| 1 | Quality : | High quality.
| 2 | Material: | Special Steel.
| 3 | Load: | (15-20 kgs cylinder).
9 |  | Oxygen Regulator and Humidifier:
| 1 | Flow Range: | From 0.15 L/min to 15 L/min.
| 2 | Maximum possible flow: | 35 L/mn above range.
| 3 | Input pressure: | Not less than 3.5 bar +/- 0,7
| 4 | Suitable for all types of fitting: | Wall outlet, trolley.
| 5 |  | Autoclavable, Unbreakable
| 6 |  | Clear Vision of Bubbles
| 7 | Humidification at 4 L/mn - 20° C: | ~3 ml/h
| 8 | Humidification at 12 L/mn - 20° C: | ~9 ml/h
| 9 | Capacity: | 100 ml (distilled water)
| 10 | Porosity of diffuser: | No More than 70 Microns
| 11 |  | Humidifier, Autoclavable, for multipurpose uses in Medical and Surgical Sections. Shall have the capability to be associated with Flowmeters (O2) or Air,O2 Reserve Cylinders and O2 Generators
| 12 |  | Fixation to flowmeters or any compatible devices shall be easy and fast (1/4 twist shall be acceptable)
10 |  | Oxygen masks (for child.&adult) & connections :
| 1 | Quality : | High quality.
| 2 | Material: | Silicone and Rubber.
11 |  | Accessories:
| 1 |  | 3 X Face Mask diffrent Size for adult. Resuable Type
| 2 |  | 3 X Face Mask diffrent Size for children. Type
| 3 |  | 1 X Tubing.
| 4 |  | 1 X Oxygen Trolley
12 |  | Estimated life-Span of the Item: | At Least 10 Years
13 |  | Warranty: | Minimum of 1 Year.
14 |  | Other Specification: | Please Specify Other Specification.
27 |  | Item Description | X-Ray Viewer
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark: | Required Certificates
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Machine should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
6 |  | Type: | Wall mounted X-Ray film viewer
7 |  | Construction: | made from acrylic translucent diffusing materials . screen thickness ≥ 3mm approx. automatic gripping mechanism included
8 |  | Functionality: | For general radiology applications. Appropriate for viewing a board band rang of film sizes
9 |  | Viewing Panel Size (L x W) , cm: | (35 x 43)
10 |  | Light: | LED
11 |  | Number of Lamps: | Please Specify
12 |  | On – Off Switch: | Included
13 |  | Power Supply: | 220 V / 50 Hz
14 |  | Estimated Life-Span of The Item: | At Least 10 Years
15 |  | Warranty: | Minimum of 1 Year.
16 |  | Other Specification: | Please Specify Other Specification.
28 |  | Item Description | Electrolyte Analyser
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall be Safe to Use.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine Should be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Method Used: | Please Specify
7 |  | Sample Type: | Whole blood, plasma, serum, dialysate, diluted urine
8 |  | Sample Volume, µL: | Please Specify
9 |  | Tube: | Open Or Closed
10 |  | Stat interrupt: | Please Specify
11 |  | Visible Sample Chamber: | Please Specify
12 |  | Analysis Time, Sec: | Please Specify
13 |  | Throughput,Samples/hr: | To facility requirements
14 |  | Assay Range, mmol/L: | Please Specify
15 |  | Test Menu:
| 1 | Na, K , CL | Included
| 2 | Sample Type: | Whole Blood /Serum/Plasma and Urine
| 3 | Other Assays: | Please Specify
16 |  | Automatic Wash: | Included
17 |  | Calibration: | Included
18 |  | Type: | Automatic
19 |  | Frequency: | Please Specify
20 |  | Reagents Close or Open System: | Please Specify
21 |  | Substitution: | Please Specify
22 |  | Display: | LED Or LCD Touchscreen
23 |  | Printer: | Built-in thermal
24 |  | Bar-Code Reader: | Please Specify
25 |  | Data Management: | Please Specify
26 |  | Interface: | Please Specify
27 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
28 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts list with code NO
29 |  | Equiped with :
| 1 |  | Solution Packs, Electrodes (Na+, K+, and CI-), conditioning Solutions, Cleaning Colutions, Calibrators & Controls
30 |  | Accessories: | Accessories and Reagents to be included in the offer
| 1 |  | Cost per Test should be Provided Clearly and Electrodes (Na+, K+, and CI-) Price
31 |  | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
32 |  | Power Supply: | AC 110-240V, 50/60Hz single-phase electrical source.
33 |  | Other Specification: | Please Specify Other Specification.
29 |  | Item Description | Ambu Bag Resuscitation Set (Infant)
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark | Required certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must Have Lattest Version
6 |  | Quality: | High Quality,
| 1 | Type: | Resuscitator, Hand-Operated, Infant (Body Weight <5 Kg), Set.
| 2 | Materials: | Latex-free
| 3 | Materials life: | All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions
7 |  |  | made of transparent silicone rubber
8 |  | Capacity: | 280 ml
9 |  | Compressible Self-Refilling Ventilation Bag: | Silicone Rubber.
10 |  | Non-rebreathing Patient Valve with Pressure Limiting Valve: | Polycarbonate/polysulfone/silicone
11 |  | Inlet Valve with Nipple for O2 Tubing: | Polycarbonate/polysulfone
12 |  | Oxygen Reservoir Bag: | Plastic.
| 1 | Masks: | Transparent silicone rubber.
| 2 | Accessories:
| 3 |  | Compressible self-refilling ventilation bag, capacity: 220-280ml
| 4 |  | Oxygen reservoir bag complete, capacity: 600ml
| 5 |  | mask 0 , face Round type, size infant 30-50 mm small
| 6 |  | mask, 1 face Round type, size infant 50-65 mm small
| 7 |  | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly
| 8 |  | Carrying case
13 |  | Estimated life-Span of the Item: | At Least 10 Years
14 |  | Warranty: | 1 Year.
15 |  | Other Specification: | Please Specify Other Specifications.
30 |  | Item Description | Ambu Bag / Resuscitation Set (Child)
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark | Required certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must Have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Type: | Resuscitator, Hand-Operated, Infant (Body Weight More Than 5 Kg), Set.
7 |  | Materials life: | All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions
8 |  |  | Made of transparent silicone rubber
9 |  | Capacity: | 700 ml
10 |  | Compressible Self-Refilling Ventilation Bag: | Silicone Rubber.
11 |  | Non-rebreathing Patient Valve with Pressure Limiting Valve: | Polycarbonate/polysulfone/silicone
12 |  | Inlet Valve with Nipple for O2 Tubing: | Polycarbonate/polysulfone
13 |  | Oxygen Reservoir Bag: | Plastic.
14 |  | Masks: | Transparent silicone rubber.
15 |  | Accessories:
| 1 |  | Compressible self-refilling ventilation bag, capacity: 500-700ml
| 2 |  | Non-rebreathing patient valve with or without pressure limiting valve, patient connector outside/inside diameter (OD/ID): 22/15mm
| 3 |  | Oxygen reservoir bag complete, capacity: 2000-2600ml
| 4 |  | mask, face shape type, size Child small
| 5 |  | mask, face shape type, size Child medium
| 6 |  | mask, face shape type, size Child large
| 7 |  | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly
| 8 |  | Carrying case
16 |  | Life Span for Item: | 7 Years
17 |  | Warranty: | 1 Year.
18 |  | Other Specification: | Please Specify Other Specification
31 |  | Item Description | Ambu Bag Resuscitation Set (Adult)
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Designed: | Heavy duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Type: | Resuscitator, hand-operated, adult (body weight over 30kg), set.
7 |  | Materials: | Latex-free
| 1 | Materials life: | All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions
| 2 |  | Made of transparent silicone rubber
| 3 | Capacity: | 2000-2600 ml
8 |  | Compressible Self-Refilling Ventilation Bag: | Silicone Rubber.
9 |  | Non-rebreathing Patient Valve with Pressure Limiting Valve: | Polycarbonate/Polysulfone/Silicone.
10 |  | Inlet Valve with Nipple for O2 Tubing: | Polycarbonate/polysulfone
11 |  | Oxygen Reservoir Bag: | Plastic.
12 |  | Masks: | Transparent silicone rubber.
13 |  | Accessories:
| 1 |  | Compressible self-refilling ventilation bag, capacity: 1475-2000ml
| 2 |  | Non-rebreathing patient valve with or without pressure limiting valve, patient connector outside/inside diameter (OD/ID): 22/15mm
| 3 |  | Oxygen reservoir bag complete, capacity: 2000-2600ml
| 4 |  | Mask, face shape type, size adult small/teenager
| 5 |  | Mask, face shape type, size adult medium
| 6 |  | Mask, face shape type, size adult large
| 7 |  | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly
| 8 |  | Carrying Case
14 |  | Life Span for Item: | 7 Years
15 |  | Warranty: | 1 Year.
16 |  | Other Specification: | Please Specify Other Specification
32 |  | Item Description | Hand Held Pulse Oximeter
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine shuld be of Latest Model and Must have Lattest Technology
| 3 | Quality: | High quality,Safety, Ease of Use
| 4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Type: | Hand Held Type
7 |  | Saturation: | 70-100 Or Better.
| 1 | Accuracy: | ±3% Or Better.
8 |  | SpO2 Range, %: | 0-100 Or Better.
| 1 | Accuracy, %: | ±3% Or Better.
9 |  | Pulse Rate bpm: | 30-240 Or Better.
| 1 | Accuracy: | ±5% Or Better.
10 |  | Response Time Sec: | Depends on application.
11 |  | Alarms :
| 1 | Audible and visual: | High/low SpO2 and pulse rate, sensor off, sensor failure, low battery
| 2 | Visual only: | High/low SpO2 and pulse rate, sensor off, low battery
12 |  | Alarm Override: | Preferred
| 1 | Reactivation method: | Automatic or manual
| 2 | Volume Control: | Required
13 |  | Self -Test Mode: | Required
14 |  | Sensor Types:
| 1 |  | Reusable Type
| 2 | Patient range: | Adult, pediatric, infant, neonate
| 3 | Cable length, m (ft): | 1.5 or More
15 | 15 | Battery:
| 1 | Type: | Rechargeable lithium-ion / or better.
| 2 | Life, hr: | 6- 12 Hr
| 3 | Recharge time, hr: | ≤3.
16 | 16 | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
17 |  | Life Span for Machine: | 10 Years
18 |  | Maintenance:
| 1 |  | 2 Years Free Maintenance.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Spare Parts list with code NO
19 |  | Required accessories:
| 1 |  | 1 X SpO2 Adult Resuable Sensor
| 2 |  | 1 X SpO2 Pediatric Resuable Sensor
| 3 |  | 1 X SpO2 Neonate Resuable Sensor
| 4 |  | 1 X Charger
| 5 |  | 1 X Carying Bag
20 |  | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
21 |  | Battery: | Rechargeable battery with built in charger with indicators for charging and low battery
22 |  | Other Specification: | Please Specify Other Specification
33 |  | Item Description | Delivery Set , Forceps
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed.
| 2 | Quality: | High Quality
6 |  | Material: | 304 Stainless Steel Or Better.
7 |  | Each Set Contents of :
| 1 |  | 1 X Knife Handle No. 3
| 2 |  | 1 X Knife Handle No. 4
| 3 |  | 5 X Backhaus Towel Clamp 5”.
| 4 |  | 2 X Rampley Sponge Holding Forceps 9.5”
| 5 |  | 1 X Mayo Scissors Straight 6 ¾”.
| 6 |  | 1 X Mayo Scissors Curved 6 ¾”.
| 7 |  | 1 X Metzembaum Scissors Curnved 6 ½”.
| 8 |  | 1 X Metzenbaum Scissors Curved 7 ½”.
| 9 |  | 1 X Episiotomy Scissors 5 ½”.
| 10 |  | 1 X Episiotomy Scissors 7”.
| 11 |  | 2 X Teals Vulsellum Forceps 9.5”.
| 12 |  | 1 X Mayo Hegar Needle Holder 6”.
| 13 |  | 1 X Mayo Hegar Needle Holder 7”.
| 14 |  | 2 X Female Catheters.
| 15 |  | 1 X Wrigley’s Outlet forceps.
| 16 |  | 1 X Dissecting forceps plain 6”.
| 17 |  | 1 X Dissecting forceps plain 7”.
| 18 |  | 1 X Dissecting forceps 1x2 teeth 7”.
| 19 |  | 1 X Instruments Box.
8 |  | Estimated life-Span of the Item: | At Least 10 Years
9 |  | Warranty: | Minimum of 1 Year.
10 |  | Other Specification: | Please Specify Other Specification
34 |  | Item Description | Ceasarian Surgical Set
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed.
| 2 | Quality: | High Quality
6 |  | Material: | 304 Stainless Steel Or Better.
7 |  | Each Set Contents of :
| 1 |  | 2 X Knife Handle Nos. 3 and 4.
| 2 |  | 6 X Backhaus Towel Clamp 5”.
| 3 |  | 2 X Rampley Sponge Holding Forceps 9 ½”.
| 4 |  | 6 X Green Armytage Forceps 8”.
| 5 |  | 2 X Dissecting Forceps Plain and tooth 6”.
| 6 |  | 2 X Dissecting Forceps Plain and tooth 8”.
| 7 |  | 2 X Metzenbaum Scissors Straight and Curved 6 ½”.
| 8 |  | 2 X Metzenbaum Scissors Straight and Curved 7 ½”.
| 9 |  | 1 X Metzenbaum Scissors Curved 8 ½”.
| 10 |  | 2 X Mayo Scissors Straight and Curved 6 ¾”.
| 11 |  | 1 X Mayo Hegar Needle Holder 6”.
| 12 |  | 2 X Mayo hegar Needle Holder 7”.
| 13 |  | 1 X Mayo Hegar Needle Holder 8”.
| 14 |  | 2 X Lengenbeck Retractor Medium.
| 15 |  | 2 X Lengenbeck Retractor Large.
| 16 |  | 1 X Doyens Retractor Medium.
| 17 |  | 4 X Allis Tissue Forceps 6”.
| 18 |  | 4 X Allis Tissue Forceps 8”.
| 19 |  | 6 X Crile Artery Forceps Straight 5 ½”.
| 20 |  | 6 X Crile Artery Forceps Curved 5 ½”.
| 21 |  | 3 X Pean Artery Forceps Straight 8”.
| 22 |  | 3 X Pean Artery Forceps Curved 8”.
| 23 |  | 2 X Teals Vulsellum Forceps.
| 24 |  | 1 X Instruments Box.
8 |  | Estimated life-Span of the Item: | At Least 10 Years
9 |  | Warranty: | Minimum of 1 Year.
10 |  | Other Specification: | Please Specify Other Specification
35 |  | Item Description | General Purpose Ultrasonic Scanner Color Doppler
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan, USA and UK
5 | 5 | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine shuld be of Latest Model and must Have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Genaral Specification: | Doppler Ultrasound is a Noninvasive Test that Can Be Used to Exam Abdomen, Abdominal Vascular, Urology, OB/GYN, Neonatal Brain, Small Parts (Breast, Thyroid, Scrotal, Prostate, MSK) , Peripheral Arterial, Peripheral Venous, Cerebrovascular.. It Works by Bouncing High-Frequency Sound Waves.
7 |  | DICOM 3.0 Compliant: | Required
8 |  | Clinical Applications : | Abdomen, Abdominal Vascular, Urology, OB/GYN, Neonatal Brain, Small Parts (Breast, Thyroid, Scrotal, Prostate, MSK) , Peripheral Arterial, Peripheral Venous, Cerebrovascular.
9 |  | Probes Types: | System Should Support 3-4 Active Probe Ports. Each Probe Must Have Several Presets with the Ability to Create User Defined Presets of Scan Settings.
| 1 | Flat Linear Array: | 5-22 MHz Required.
| 2 | Convex Array: | 2.5-6 MHz Required.
| 3 | Endovaginal: | 5-9 MHz Required.
| 4 | Dedicated 3-D/4-D: | Required.
10 |  | Scan Mode :
| 1 | 3-D (Freehand): | Required
| 2 | 3-D (Automatic): | Required
| 3 | 4-D (Real-Time 3-D): | Required
| 4 | B-Mode (2-D): | Required
| 5 | Contrast Enhanced Ultrasound (CEUS). | Required
| 6 | Tissue Harmonic B-mode: | Required
| 7 | M-Mode: | Required
| 8 | Doppler: | Required
| 9 | Color Doppler Imaging (CDI) 3-D/4-D: | Required
| 10 | Power Doppler imaging (PDI) 3-D/4-D: | Required
| 11 | Continuous Wave: | Required
| 12 | Pulsed Wave: | Required
| 13 | Duplex: | Required
| 14 | Triplex: | Required
| 15 | Tissue: | Required
| 16 | Elastography: | Required
| 17 | Strain (Manual): | Required
| 18 | 2-D Shear Wave: | Required
| 19 | Point Shear Wave: | Required
| 20 | Quantitative: | Required
| 21 | Needle Enhancement Mode/Capability: | Required
| 22 | Others: | Biopsy Guidance Required
11 |  | Image Display and Processing:
| 1 | Automated B-Mode Image Optimization: | Required
| 2 | Grayscale Levels: | ≥256 Or Better.
| 3 | Image Magnification (Zoom): | Required
| 4 | Monitor cm (in) : | Split screen (Twin View)
| 5 |  | Large Screen 20 -23 inch (at least) color LCD/LED monitor with swivel arm to allow rotation/ movement in/out and up/down. And 12-15 Inch for Small Screen
| 6 | Maximum Display Depth, mm: | ≥240 Or Better.
| 7 | Preprocessing | Required
| 8 | Postprocessing | Required
12 |  | Digital Image Storge:
| 1 | Cine Loop Playback: | Required
| 2 | Data Storage and Backup: | Digital storage hard drive, ≥500 GB Or Better
13 | 13 | Software Features:
| 1 | Operating System: | Please Specifiy
| 2 | Analysis Packages: | Multiple Including , Abdominal, Cardiac, Vascular, OB/GYN, MSK, Small Parts, Thyroid, Urology , Liver , Breast , Spine and Others Parts.
| 3 | Digital Calipers: | Required
| 4 | Exam Presets: | Required
| 5 | On-Screen Annotation: | Required
14 |  | Connectivity:
| 1 | Data Export Methods: | Ethernet and Wirless Preferred.
| 2 | Exported Still Image Formats: | JPEG, DICOM
| 3 | Exported Video Clip Formats: | Required
15 |  | Reporting: | The system should have an integrated reporting software with the ability to customize the design, add the logo and etc.
| 1 |  | Reports Must include Exam results including patient info, exam info, measurements, calculations, images, comments and diagnosis of the doctor.
| 2 |  | Several Standard templates should be provided and able to add new report templates as required.
| 3 |  | The software should be able to convert reports to PDF formats.
| 4 |  | Reports must be printable to any regular office color printer with high quality settings.
16 |  | User Interface: | The system should include full alphanumeric keyboard.
| 1 |  | Touch screen interface is very much recommended.
| 2 |  | Ultrasound functions keyboard should have a presets buttons for easy access of certain functions in the software.
| 3 |  | System should have at least 6 TGC Pods, with Re-mapping Functionality at Any Depth.
17 |  | User Documentation: | On board User Manual.
| 1 |  | Service Manual
| 2 |  | Printed Hard Copy of user Manual Must Be attached with each system.
18 |  | Accessories:
| 1 |  | UPS Online 2 KVA
| 2 |  | B\W Video Printer
| 3 |  | Color Printer Laser Type.
| 4 |  | 6 X Gel, 500 g
| 5 |  | 6 X Roll Paper
| 6 |  | 1 X Dust Cover.
19 |  | Essential Requirement:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years or more.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
20 |  | Maintenance:
| 1 |  | Preferred Less Maintenance Needed. 3 Years Free Maintenace, Including PM Kit.
| 2 |  | Service Manual Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts List With code No.
| 5 |  | Including Maintenance and calibration Tools.
21 |  | Life Span for Machine: | At Least10 Years.
22 |  | Training: | Training (2) Doctors to Operate and use the Device, and Training (1) Medical Device Engineers to Maintain and Calibrate the Device (Full Training Program)
23 |  | Self Diagnostic Error / Calibration Massage: | Included
24 |  | Power Supply: | 100-240 V, 50/60 Hz Electriv Saftey Class I.
25 |  | Other Specification: | Please Specify Other Specification
36 |  | Item Description | Suture Instruments and Dressing Set
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed.
| 2 | Quality: | High Quality
6 |  | Material: | 304 Stainless Steel Or Better.
7 |  | Each Set Contents of :
| 1 |  | 1 X Tray Instrument with Lid 225X 125 X 50 mm Deep
| 2 |  | 6 X Forceps artery Kocher COF BJ.S/S 180mm.
| 3 |  | 1 X Forceps Dissecting Bonney's 1X 2 Teeth S/S 150 mm
| 4 |  | 1 X Forceps Dissecting Plain serrated jaws S/S 150 mm
| 5 |  | 2 X Forceps Dressing Bryants serrated jaws S/S 125mm
| 6 |  | 4 X Forceps Towel Holding cross Action S/S 100mm
| 7 |  | 2 X Forceps Sponge Holding Rampley BJ S/S 240 mm
| 8 |  | 1 X Handle size 3 for Scalple blades size 10-15
| 9 |  | 1 X Needle Holder Kilner's 165 mm
| 10 |  | 1 X Scissors Mayo Straight SJ. S/S 200mm.
| 11 |  | 1 X Scissors Iris straight, Sharp points 115 mm
| 12 |  | 1 X Sterilizing and storing case St. St. with filters 230x130x50 mm
8 |  | Estimated life-Span of the Item: | At Least 10 Years
9 |  | Warranty: | Minimum of 1 Year.
10 |  | Other Specification: | Please Specify Other Specification.
37 |  | Item Description | Large Dressing Set
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed.
| 2 | Quality: | High Quality
6 |  | Material: | 304 Stainless Steel Or Better.
7 |  | Each Set Contents of :
| 1 |  | 1 X Tray Instrument with Lid 225X 125 X 50 mm Deep
| 2 |  | 6 X Forceps artery Kocher COF BJ.S/S 180mm.
| 3 |  | 1 X Forceps Dissecting Bonney's 1X 2 Teeth S/S 150 mm
| 4 |  | 1 X Forceps Dissecting Plain serrated jaws S/S 150 mm
| 5 |  | 2 X Forceps Dressing Bryants serrated jaws S/S 125mm
| 6 |  | 4 X Forceps Towel Holding cross Action S/S 100mm
| 7 |  | 2 X Forceps Sponge Holding Rampley BJ S/S 240 mm
| 8 |  | 1 X Handle size 3 for Scalple blades size 10-15
| 9 |  | 1 X Needle Holder Kilner's 165 mm
| 10 |  | 1 X Scissors Mayo Straight SJ. S/S 200mm.
| 11 |  | 1 X Scissors Iris straight, Sharp points 115 mm
| 12 |  | 1 X Sterilizing and storing case St. St. with filters 230x130x50 mm
8 |  | Estimated life-Span of the Item: | At Least 10 Years
9 |  | Warranty: | Minimum of 1 Year.
10 |  | Other Specification: | Please Specify Other Specification.
38 |  | Item Description | Dressing Trolley
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates.
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan,USA and UK
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
| 2 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Material of Construction: | Heavy duty Stainless Steel 304 or better and high quality
7 |  | General Description: | Dressing trolley with two non-removable shelves mounted on four swivel castors.
8 |  | Technical Specifications: | Guard rails on three sides on top shelf and Second Shelf
| 1 |  | 2 level shelves
| 2 | Carrying Capcity Approx: | 80-100 Kg
| 3 | Mobile On Swivel Castors : | 100mm (2 with brakes).
9 |  | Accessories:
| 1 |  | 1 x Dressing Drum Holder.
| 2 |  | 1 x Thermometer Holder.
| 3 |  | 1 x Forceps Container Holder.
| 4 |  | 1 x S.S. Dressing tray attached at the top. With Collapsible Holder
| 5 |  | 1 x Waste bin
| 6 |  | 1 x S.S. drawer below the first shelf.
| 7 |  | 1 x S.S Sponge Bowl.
| 8 |  | 1 x Dressing Jar.
10 |  | Guiding Dimentions of trolley : | 650-800Lx400-500 Wx900-1000mmH at least
11 |  | Life Span for Item: | 7 Years
12 |  | Warranty: | Minimum of 2 Years
13 |  | Other Specification: | Please Specify Other Specification
39 |  | Item Description | Stand Drip IV
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Type: | Drip Stand or Rod for I.V. Procedures
7 |  | Construction: | Rod made of chrome-plated steel tubing, Adjustable heightonstructed from heavy duty section steel tube
8 |  |  | Shall be mobile on very stable stand with 5 castors, 5 stars-shaped. Castors shall be orientable and shall have a smooth running
9 |  |  | Rod equipped with 4 hooks configuration.
10 |  | Estimated life-Span of the Item: | At Least 7 Years
11 |  | Warranty: | Minimum of 2 Years.
12 |  | Other Specification: | Please Specify Other Specification
40 |  | Item Description | Ward Screen 4 Sections
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
4.1 | 1 | FDA Certificate Or CE Mark | Required Certificates
4.2 | 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, USA and UK
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must Have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Design & Quality: | Bed screen, hospital, on castors.
7 |  | Construction: | Metal construction,painted with elecrostatic powder,4 screens / 4 folds complete with Pvc, mobile castors.
8 |  | General Description: | Bed screen, hospital, on castors.
9 |  | Technical Specifications:
| 1 |  | Hospital bed screen in 4 sections with curtains
| 2 |  | Bed screen mounted on swivel castors; positioned in broad stance for good stability.
10 |  | Frame: | Epoxy or enamel coated or chrome plated tubular steel.
11 |  | Curtains: | Plastic non-transparent, washable and flame retardant.
12 |  |  | Extendable / Retractable
13 |  |  | Multiple leg supports for stability.
14 |  |  | Hygienic, easy to clean panels.
15 |  | Castors: | Mobile on 50-75mm dia. swivel castors
16 |  |  | complete with pvc curtain.
17 |  | Guiding Dimentions: | 180-200 cm(h) x 200-220 cm(w) x 45-50 cm(d)
18 |  | Estimated life-Span of the Item: | At Least 7 Years
19 |  | Warranty: | Minimum of 2 Years.
20 |  | Other Specification: | Please Specify Other Specification
41 |  | Item Description | Blood Bank Plasma Deep Freezer (-40 °C)
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Machine Should be of latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality,Safety, Ease of Use
| 4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The company.
6 |  | Type: | Approved standard electrical Blood Bank plasma freezer that uses a compressor circulating CFC-free refrigerant.
| 1 |  | Upright type
7 |  | Purpose of Equipment: :
| 1 |  | To Freeze or store Plasma.
| 2 |  | Must be designed specifically for blood bank use. Commercial or modified commercial freezers for other purpose are not acceptable.
8 |  | Technical Specifications :
| 1 | Temperature Range: | -20° C to ­40° C
| 2 | Capacity : | 200 Standard plasma bags. Or More
9 |  | Material Construction:
| 1 | Inner Chamber: | Made of Stainless Steel SS­304, 22 SWG
| 2 | Outer Chamber : | Made of Powder Coated CRCA sheet, 18 SWG
| 3 | Powder Coating Thickness: | Insulation polyurethane foam >80mm thick, foaming agent CFC free.
| 4 | Sub­Doors : | Made of Stainless Steel
| 5 |  | Should be mounted on lockable caster wheels
10 |  | Door
| 1 |  | Separate inner doors to prevent cold loss
| 2 |  | Automatic/Magnetic closing of at least inner doors.
| 3 |  | Heating device in front to avoid condensation
| 4 |  | Opening angle limited (eg <135°) Or Better.
| 5 |  | Door open/ajar audio and visual alarm
| 6 |  | Door lock should be available
11 |  | Shelves / Trays: | 4 Nos Or more. of stainless steel shelves made of SS­304 (22 SWG) and 4 compartments
12 |  | Hot Line: | Anti­moisture entry using hot line at the mouth of the cabine
13 |  | Compressors: | Hermetically sealed compressors
14 |  | Control System: | Microprocessor Control & LCD display .
15 |  | Internal Temperature:
| 1 |  | Should be able to maintain internal temperature not warmer than -30°C to -35°C.
| 2 |  | Whatever the load, setting accuracy less than or equal to 1°C
| 3 |  | Temperature should not go Outside Safe Range.
16 |  | Hold-Over Time:
| 1 |  | A full load of plasma packs at -36 °C takes at least 1 hr to rise to above -20 °C
| 2 |  | A full load of plasma packs at -36 °C takes at least 32 hrs to rise to above -5 °C
17 |  | Cooling Down Time:
| 1 |  | A full load of plasma packs at +25°C takes a maximum of 5 hrs for all the packs to reach below -5 °C
| 2 |  | A full load of plasma packs at +25 °C takes a maximum of 30 hrs for all the packs to reach below -20 °C
18 |  | Temperature monitoring, Thermograph and Related Alarms:
| 1 |  | Digital temperature (LED) display with 0.1 °C graduation.
| 2 |  | Microprocessor controlled primary temperature control
| 3 |  | Integrated Visual AND Audible Temperature alarm systems,
| 4 |  | There should be a method to test the alarm system
| 5 | Alarm History: | Temperature maximum and minimum, average temperature during alarm period, time of duration of alarm.
| 6 |  | Provision to be connected to a remote monitoring system and remote alarm.
| 7 |  | The temperature record should be electronically logged (that can be retrieved eg by USB port) and also documented on a physical thermograph; preferably with a 7-day graphic chart recorder with supply of free charts for full period of warranty.
| 8 |  | Must have Battery backup for temperature recordings which is especially needed during power failure/fluctuations
| 9 |  | Additional Battery backup for alarm so that alarm will not fail in case of power failure, and should be able to sustain the alarm.
19 |  | Alarms: | High and low temperature alarm , Sensor Failure alarm , Temperature controllers Failure alarm .
20 |  | Refrigeration Type: | Direct Refrigeration.
21 |  | Defrosting: | Automatic / Manual .
22 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
23 |  | Life Span for Machine: | 10 Years
24 |  | Maintenance:
| 1 |  | 2 years free maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
25 |  | Accessories:
| 1 |  | Fuses Set
| 2 |  | Filters Set
| 3 |  | Supplied with Complete Accessories .
26 |  | Estimated life-Span of the Laboratory Deep Freezer: | At Least 10 years Or More.
27 |  | Power Supplay: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase ( power cable Compatible with the Hospital electric outlet plug, 5 mt ), Electrical Safety class 1
28 |  | Other Specification: | Please Specify Other Specification
42 |  | Item Description | Platelets Incubator with Agitator
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy duty Designed, New Model.
| 2 |  | The Machine Should be of latest Model and Must have Lattest Technology
| 3 | Quality: | High Quality,Safety, Ease of Use
| 4 | Supplier | The Supplier is Required to be a Certified Distributor or Agent of The company.
6 |  | Specifications:
| 1 |  | Micro processed controlled LED/LCD display, temperature graph display & digital sensor probe with Audio Visual Alarm for high & Low Temperature & Door Open
| 2 |  | Flatbed agitator fitted inside a temperature controlled incubator that uses CFC-free refrigerant and CFC free insulation material .
| 3 | Capacity: | At least 40-50 standard random platelet unit bag
| 4 |  | The unit is mounted on castor wheel for easy mobility.
| 5 |  | Should be equiped with CFC Free PUF (Poly Urethane Foam) Insulation .
| 6 | The Operating Temperature: | 22 °C± 0.5 -1.0 °C .
| 7 | The Place for Agitator: | 2 Nos .Or More
| 8 |  | Extra-Strength Non CFC Refrigeration System .
| 9 |  | Drawers Perforated to Ensure Good Air Circulation .
| 10 |  | The Agitator Holding The Shelves is Suspended in Such a Way as to Ensure Minimum Noise for The Life of The Agitator.
| 11 |  | Glass door with Full Visibility of Units Without Opening Door .
7 |  | Alarm | Digital sensor probe with Audio Visual Alarm for:
| 1 |  | High & Low Temperature
| 2 |  | Door Open
| 3 |  | Agitator Motion Fault
8 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 5 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
9 |  | Life Span for Machine: | 10 Years
10 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
11 |  | Training:
| 1 |  | User /Nurses Training, By Specialist From the Supplier.
12 |  | Guarantee:
| 1 |  | The Vendor Should Guarantee The Service and Spare Support for 10 Years of The System and all Accessories after 2 Years of Warranty
| 13 | Power Requirments: | Power Input to be 220-240VAC, 50/60Hz
| 14 | Other Specification: | Please Specify Other Specification.
43 |  | Item Description | Stethoscope Binaural Bowels Stainless Steel
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK.
5 | 5 | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Type: | Stethoscope Binaural Bowels
7 |  | Quality: | Heavy Duty and High Quality
8 |  | Product Description:
| 1 |  | A Mechanical listening device designed for listening to sounds from the heart and lungs. Comprises a membrane atthe listening head connected by a split "Y" tube to the headgear with ear pieces that are placed into the users ears.
9 |  | Specifications:
| 1 |  | Double cup, dual-use (adult and pediatric auscultation) chest piece in stainless steel or chrome plated brass.
| 2 |  | Adult diaphragm 43-47mm; pediatric diaphragm 28-36mm.
| 3 |  | Y tube treated rubber or PVC with 8-11mm diameter.
| 4 |  | Sensitivity from 3.2dB to 26dB in a range from 50 to 1000Hz for cardiology.
| 5 |  | Sensitivity 8.1dB in a range from 600 Hz to 1,500Hz for pneumology.
| 6 | Arms: | Stainless Steel
| 7 |  | Removable Plastic Ear-Pieces.
| 8 |  | Latex-free
10 |  | Accessories:
| 1 |  | 4 X Ear caps made of elastomer rubber
| 2 |  | 1 X Storge Case
| 3 |  | 1 X Operation Manual
11 |  | Estimated life-Span of the Item: | At Least 10 Years
12 |  | Warranty: | Minimum of 1 Year.
14 |  | Other Specification: | Please Specify Other Specifications.
44 |  | Item Description | Thermometer Digital
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates.
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 |  | Designed: | Heavy duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Power Type: | DC Complete.
7 |  | Quality: | Heavy duty and high quality
8 |  |  | For oral, underarm, or rectal use
9 |  | Operation Details: | Press ON/OFF button to activate. The unit would beep and display.
10 |  | Display: | This LCD display test will go on for around 2 seconds
11 |  | Low Temperature Display: | Temperature < 32°C Display L°C.
12 |  | High Temperature Display: | Temperature > 43°C Display H°C.
13 |  | Accuracy: | 35°C ----- 39°C ±0.1°C.
14 |  | Display Resolution: | 0.1°C.
15 |  | Normal Beep Sound: | 10 long beep sounds in 10 seconds.
16 |  | Fever Beep Sound: | (>37.8°C) : 30 short beep sounds in 10 seconds.
17 |  | Auto Power Off: | 8 minute ±2 minute
18 |  | Battery: | AG3 or LR41
19 |  | Life Of Battery: | 2 years or 2500 times.at least
20 |  | Accessories:
| 1 |  | 1 X Plastic Case.
| 2 |  | 2 X Spare Part Battery.
| 3 |  | Supplied With Complet Accessories.
21 |  | Life Span for Item: | 7 Years
22 |  | Warranty: | Minimum of 1 Year.
23 |  | Other Specification: | Please Specify Other Specifications.
45 |  | Item Description | Thyroid Surgery Instruments Set
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed.
| 2 | Quality: | High Quality
6 |  | Material: | 304 Stainless Steel Or Better.
7 |  | Each Set Contents of :
| 1 |  | 2 X Kocher Thyroid Dissectors.
| 2 |  | 2 X Joll’s Thyroid Self-Retaining Retractor.
| 3 |  | 2 X Lengenbeck Retractor Medium.
| 4 |  | 2 X Kilner Catspar Retractor.
| 5 |  | 2 X Single Hook Retractor.
| 6 |  | 1 X Dissecting forceps 1x2 teeth 6”.
| 7 |  | 1 X Dissecting forceps plain 6”.
| 8 |  | 1 X Metzenbaum Scissors Curved 6 ½”.
| 9 |  | 1 X Metzenbaum Scissors Curved 7 ½”.
| 10 |  | 1 X Knife Handle No. 3
| 11 |  | 1 X Knife Handle No. 4.
| 12 |  | 5 X Backhaus Towel Clamp 5”.
| 13 |  | 2 X Rampley Sponge Holding Forceps 9.5”
| 14 |  | 2 X Lahey Thyroid Vulsellum Forceps 6”.
| 15 |  | 1 X Instruments Box.
8 |  | Estimated life-Span of the Item: | At Least 10 Years
9 |  | Warranty: | Minimum of 1 Year.
10 |  | Other Specification: | Please Specify Other Specification
46 |  | Item Description | Scale for Adult with Height Measurement
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates.
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
| 2 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Type: | Scale for Adult with height measurement
7 |  | Design: | Heavy duty and high quality.
8 |  | Capacity: | >200 Kg
9 |  | Display: | LCD
10 |  | Auto Zero Switch: | Recuired
11 |  | Auto Calibration: | Recuired
12 |  | Battery: | Re-Chargeable Battery
13 |  | Power supplay: | 100 -240 V 50Hz, Electric adapter
14 |  | Life Span for Item: | 7 Years
15 |  | Warranty: | Minimum of 1 Year.
16 |  | Other Specification: | Please Specify Other Specification.
47 |  | Item Description | Lamp, Examination Mobile Complete /Emergency
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates.
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, USA and UK.
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  | Design: | Mobile
| 1 | Type: | Lamp, Examination. Mobile on 4\5 castors, 5 stars-shaped heavy and stable stand.
| 2 | Quality: | Heavy duty and High quality
7 |  | Lamp Type: | Led's type
8 |  |  | Height of Lamp head shall be adjustable
9 |  | Light Intensity: | Shall be at least 25000-40000 Lux at 50 cm.min. with a fields radius of 17cm minimum
10 |  | Color Intensity: | no less than 4300 °K. Single Aluminum reflector of 150mm minimum
11 |  | Power Rating: | 1 X 3W Minimum. 220V/50Hz
| 12 | Power Supply: | 220V/50 Hz. Transformer, 12v-24V DC shall be included must be in conformity with Council Direct 93/42/EEC, on medical devices
| 13 | Battery Emergecy : | Bakup 2 hr
| 14 | Accessories:
| 14.1 |  | 3 Fuses
| 15 | Life Span for Item: | 7 Years
| 16 | Warranty: | Minimum of 2 Years.
| 17 | Other Specification: | Please Specify Other Specification.
48 |  | Item Description | Torch Pen Light For Examination
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates.
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 |  | Designed: | Heavy duty Designed, New Model.
| 1 | Quality: | High Quality.
| 2 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Power Type: | DC Complete.
7 |  |  | Torch Pen Light For Examination, Spot Light
8 |  |  | High-performance pen-lights for medical diagnosis.
9 |  |  | Available with LED illumination.
10 |  |  | Bright and natural white light by high-performance LED
11 |  | LED Life | at least 20.000 hours
12 |  | Material: | ABS casing with metal clip, disinfectable
13 |  |  | Easy to use with on/off switch
14 |  | Accessories:
| 1 |  | 1 X Case
| 2 |  | 3 X Spare Battery
| 3 |  | 2 X Spare Led.
15 |  | Life Span for Item: | 7 Years
16 |  | Warranty: | Minimum of 1 Year.
17 |  | Other Specification: | Please Specify Other Specifications.
49 |  | Item Description | Oxygen Two Stage Regulators
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | CE Mark | Required Certificates
5 |  | Designed: | Heavy Duty Designed, New Model.
| 1 | Quality: | High Quality.
6 |  |  | Oxygen Two Stage Regulator , The Pressure Rate of Medical gas Supply From Cylinder to Medical Device.
| 1 | Construction | Chrome Plated Brass Body , Steel or Equivalent.
| 2 |  | Double Chamber Gas Control
| 3 |  | Fits Cylinders with Vertical Outlet Valves
| 4 |  | Two Gauges Display- Bottle Contents and Output Pressure
| 5 |  | Pressure Controlled by The Centre Dial
| 6 | Maximum Outlet Pressure | Not less than 10 Bar
| 7 | Maximum Inlet Pressure | Not less than 300 Bar
| 8 |  | Autoclavable, Unbreakable
| 9 |  | Clear Vision of Bubbles.
7 |  | Estimated life-Span of the Item: | At Least 10 Years
8 |  | Warranty: | Minimum of 1 Year.
9 |  | Other Specification: | Please Specify Other Specification.
50 |  | Item Description | Tonsillectomy Instruments Set
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed.
| 2 | Quality: | High Quality
6 |  | Material: | 304 Stainless Steel Or Better.
7 |  | Each Set Contents of :
| 1 |  | 1 X Davis - Boyle Mouth Gag Set.
| 2 |  | 2 X Suspension Rods.
| 3 |  | 1 X Suspension Plate.
| 4 |  | 1 X Gwyneven Tonsil Dissector.
| 5 |  | 1 X Mollison Anterior Pillar Retractor.
| 6 |  | 1 X Tonsil Dissector with Sucker.
| 7 |  | 1 X Waugh Tonsil Dissecting Forceps serrated 8”.
| 8 |  | 1 X Waugh Tonsil Dissecting Forceps 1x2 teeth 8”.
| 9 |  | 1 X Brikett Tonsil Artery Forceps Fine Straight.
| 10 |  | 1 X Negus Tonsil Artery Forceps Short Curve.
| 11 |  | 1 X Negus Tonsil Artery Forceps Large Curve.
| 12 |  | 1 X Wilson Tonsil Artery Forceps.
| 13 |  | 1 X Metzenbaum Tonsil Scissors Straight 7 ½”.
| 14 |  | 1 X Metzenbaum Tonsil Scissors Curved 7 ½”.
| 15 |  | 1 X McIndoe Tonsil Scissors Fine Point Curved 7 ½”.
| 16 |  | 1 X Eve’s Tonsil Snare.
| 17 |  | 1 X Tonsil Snare Wire Packet. (Packet of 15 Pcs.)
| 18 |  | 1 X St. Clair Thomson Adenoid Curette with Guard Small.
| 19 |  | 1 X St. Clair Thomson Adenoid Curette with Guard Large.
| 20 |  | 1 X Negus Pusher.
| 21 |  | 1 X Denis Brown Tonsil Holding Forceps.
| 22 |  | 1 X B.P. Knife Handle No. 7.
| 23 |  | 1 X Yankauer Suction Tube.
| 24 |  | 1 X Instruments Box size.
8 |  | Estimated life-Span of the Item: | At Least 10 Years
9 |  | Warranty: | Minimum of 1 Year.
10 |  | Other Specification: | Please Specify Other Specification
51 |  | Item Description | Vaginal Speculum Set
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed.
|  | Quality: | High Quality
6 |  | Material: | 304 Stainless Seel Or Better
7 |  | Each Set Contents of :
| 7.1 |  | 2 X Speculum Vaginal Cuscos - Medium
| 7.2 |  | 2 X Speculum Vaginal Cuscos - Small
| 7.3 |  | 2 X Speculum Vaginal Cuscos - Large
8 |  | Estimated life-Span of the Item: | At Least 10 Years
9 |  | Warranty: | Minimum of 1 Year.
10 |  | Other Specification: | Please Specify Other Specification.
52 |  | Item Description | Patient Stretcher
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan,USA and UK
5 | 5 | Genaral Specification: | The Unit Shall Be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Item Must have Lattest Version
| 3 | Quality: | High Quality,
6 |  | Construction & Features: | Constructed from steel frame epoxy powder coated painting
| 1 | Lever Operated: | back raise
| 2 | Bed side Rails: | Sliding aluminum or chrome plated steel
| 3 | Mattress: | 7cm urethane foam approx.
| 4 | Bumpers: | Perimeter
| 5 |  | Trendelenburg position
| 6 | IV Pole Sockets, Number: | ≥4, at least 1 mountable at each corner
| 7 |  | Crank-operated high-low position
| 8 | Oxygen Tank Holder: | Required
7 |  | Casters:
| 1 | Diameter, cm: | ≥12.7, 20.3 preferred
| 2 | Brakes: | ≥2, diagonal
8 |  | Maximum Patient Wehgth, kg: | ≥200 kg
9 |  | Stretcher Type: | Adjustable height
10 |  | Dimension:
| 1 | (L x W),cm: | (210 x 80) or better.
| 2 | Height, cm: | Height adjustment by hedrulic system (50 – 120)
11 |  | Estimated life-Span of the Item: | At Least 7 Years
12 |  | Warranty: | Minimum of 2 Years.
13 |  | Other Specification: | Please Specify Other Specifications.
53 |  | Item Description | Semi Automated Biochemistry Analyzer
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan, USA and UK
5 | 5 | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine shuld be of Latest Model and must Have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | Technical Specification : | Should be microprocessor controlled general purpose bi-chromatic photometer
| 1 | Filters: | 6-7 Filters.
| 2 | Ranging nm: | 340 to 660 nm Or Better.
| 3 | Light source : | Led Lamp
| 4 | Methods: | End point, kinetic and two point kinetic measurement modes or more
7 |  | Printer: | Build-in
8 |  | Measurement Range : | from 0.001 to 2.300Abs or better
9 |  | Display: | LCD display. Touch Screen
10 |  | Control: | Quality control – two control/test QC survey of at least 30 points, Levy Jenny plot.
11 |  | Bandwidth nm: | 5 nm or lesser.
12 |  | Programme Memory: | Should have a test programme memory of 50 or more.
13 |  |  | Should be provided with sample carry over prevention facility.
14 |  |  | Aspiration should be based on Bellow/Peristaltic Pump/ Vacuum pump.
15 |  | Reagents : | Should provide: 500 ml of reagents for urea, S. creatine, S. bilirubin, sugar, cholesterol, and Quality control 5ml one each for normal, abnormal.
16 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
| 5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
| 7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
| 8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
| 9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
17 |  | Life Span for Machine: | 10 Years
18 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
| 4 |  | Spare Parts list with code NO
19 |  | Accessories:
| 1 | Bulb: | One Spare bulb
| 2 | Fues: | 3 Fueses
20 |  | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
21 |  | Power Supplay: | 100 to 240 V ~ ±10%, 50/60 Hz Single Phase automatic range selection ( power cable Compatible with the Hospital electric outlet plug, 5 mt ), Electrical Safety class 1
22 |  | Other Specification: | Please Specify Other Specification.
54 |  | Item Description | Blood Bank Refrigerated Centrifuge
No. | Sub No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 |  | Manufacturer : | Please Specify Manufacturer Name.
2 |  | Model Number: | Please Specify Model Number of The Offered Equipment.
3 |  | Country Of Origin: | Please Specify Country of Origin of The Offered Equipment.
4 |  | Safety Standard:
| 1 | FDA Certificate Or CE Mark: | Required Certificates
| 2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union , Japan, USA and UK
5 |  | Genaral Specification: | The Unit Shall be Safe to Use
| 1 | Designed: | Heavy Duty Designed, New Model.
| 2 |  | The Machine shuld be of Latest Model and must Have Lattest Technology
| 3 | Quality: | High Quality, Safety, Ease of Use
| 4 | Supplier: | The Supplier is Required to be a Certified Distributor or Agent of The Company.
6 |  | General Description: | Blood Bank Refrigerated centrifuge is a mechanical device for separation of blood components like packed red cells, platelet rich plasma, platelet concentrates ,plasma, cryoprecipitate & Buffy Coat.
7 |  | Control: | It should have a Programmable microprocessor control system with self-diagnostic feature and Microprocessor Control operation with automatic calculation of RPM to RCF.
8 |  | Operation: | Unlimited and defined programming of all parameters, automatic programming sequence controlled by microprocessor as well as.
9 |  | Display: | Digital Display of Nominal and Actual Values of All Parameter Indication (Time, Speed, Temp, Rotor Recognition, etc.)
| 1 |  | User-friendly Equipment- Easy to read digital display for controlling basic Functions.
10 |  |  | Auto-door function to open and close the centrifuge door automatically using simple push-button and instant rotor opening and closing with lid storage.
11 |  | Safety: | There should be a Lid lock and interlock- Automatic Double locking for operator safety. The centrifuge lid locking should be automatic and motorized. No manual push or banging of lid should be required. Door of the centrifuge will not open during power failure by pressing of regular door open switch or Knob or button for Operator Safety.
12 |  | Max Capacity: | 4 x 1000 ml Or Better With Double Blood Bag Buckets
| 1 | Maximum Speed: | Above 4500 -5000 RPM Or Better.
| 2 | Maximum RCF : | Please Specify
13 |  |  | It should have Facility to remove the blood bags during power failure/emergency by opening centrifuge door through mechanical override release cords located on both sides of centrifuge. This should be known to trained persons only for intentional use. To protect blood bags and operator, the door should not open with regular electrical knob/button/switch during a power failure.
14 |  | Stainless Steel Chamber: | Easy to Clean, Corrosion Resistant with Provision of Both Drain and Condensed Water Collection Container for Complete Decontamination incase of blood bag rupture/leakage.
15 |  | Run Time: | 99 hours 59 sec (1 second increment) .
16 |  | Program Memory: | Capacity to store at least 100 Centrifugation Programs Or More.
17 |  | Temperature Control Range: | (-20°C to 40°C, adjustable to 1° C. )
18 |  |  | CFC free refrigeration.
19 |  | Acceleration Rates: | Please Specify
20 |  |  | Audio-Visual Alarms if: lid open, operation end, brake, over temperature and imbalance detection.
21 |  | Accessories:
| 1 |  | To be Supplied with Blood Bag Rotor with Double Blood Bag Buckets. (To Double the Capacity).
| 2 |  | With All Accessories
22 |  | Certification from the Manufacturer:
| 1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
| 2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
| 3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
| 4 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
| 5 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
23 |  | Life Span for Machine: | 10 Years
24 |  | Maintenance:
| 1 |  | 2 Years Free Maintenace.
| 2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
| 3 |  | Application Software and Interface Connection Included.
25 | 25 | Training:
| 1 |  | User /Nurses Training, By Specialist From the Supplier.
26 |  | Guarantee:
| 1 |  | The Vendor Should Guarantee The Service and Spare Support for 10 Years of The System and all Accessories after 2 Years of Warranty
27 |  | Power Supply: | 100 to 240 V ±10%, 50 Hz.
| 1 | Power Consumption: | Please Specify
28 |  | Other Specification: | Please Specify Other Specification.
Company Name اسم الشركة |  |  |  | Signature & Stamp التوقيع و الختم
Representative Nameاسم ممثل الشركة
Designation المسمى الوظيفي
Date: التاريخ

## Technical Specification(LOT2) (hidden)
Annex B (LOT 2)- Technical Specification نموذج المواصفات الفنيه Equiping Operation Theater for Aslam HC and Al Shaheed Al Kuhlani Hospital - Hajjaah
Technical Specifications
OT-1 | Item Description | Anesthesia Machine with Ventilator
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the company.
6 | System Components:
6.1 |  | Original cart with antistatic braked castors, shelves and 2-3 drawers.The basic unit must form rigid stable and easy for use.
6.2 |  | Ventilator
6.3 |  | Multi vaporize capability
6.4 |  | Regulators and flow meters
6.5 |  | CO2 absorber
6.6 |  | Patient type: adult, Pediatric.Neonat
6.7 |  | patient monitor
7 | Circulation system:
7.1 |  | Compact
7.2 |  | Sterilizable
7.3 |  | Changeable from rebreathing (closed) to non-rebreathing (open)
8 | Gas mixing system:
8.1 |  | Mechanical safety system
8.2 |  | O2 flow meters (One tube)
8.3 |  | N2O flow meters (One tube)
8.4 |  | Air flow meters (One tube)
8.5 |  | Nitrous oxide blockage with constant alarm
8.6 |  | N2O , O2 ratio control
8.7 |  | Control of gas supply for O2 and air and for all gases
8.8 |  | Safety valve built in to patient system
9 | CO2 absorber:
9.1 |  | Bag /ventilator switch
9.2 |  | By-pass switch
9.3 |  | Pressure reading (Manometer or Electronic)
9.4 |  | canester ( 2 kg capacity or 1.5 L)
10 | Vaporizer:
10.1 |  | Selecta tic type
10.2 |  | Two vaporizers
10.3 |  | Fast coupling system
10.4 |  | Supplied with sevoflurane
10.5 |  | Supplied with isoflurane
11 | Gas Anesthesia Monetoring: | Included
12 | O2/N2O Cylinders Yokes with Cylinders: | Included
13 | O2–flush: | Included
14 | Mechanical Anti Hypoxic Device (AHD): | Included
15 | Central Gas Supply: | O2/N2O/Air, medical quality 3-6 bar
| (Pipeline Gas Inlet): | Selector switch for O2/N2O and O2/Air
16 | Flow Meters:
16.1 |  | Color coded
16.2 |  | O2 : 0.05-10 L/min or better
16.3 |  | Air: 0.2-12 L/min or better
16.4 |  | N2O: 0.05-10 L/min or better
17 | Ventilator | Electronically controlled
17.1 | Bellow system | Range from pediatric to adults
Standing , easily seen
17.2 | Ventilation modes (selectable)
17.2-1 |  | Stand by
17.2-2 |  | Spontaneous
17.2-3 |  | Manual
17.2-4 |  | CMV (Pressure control/ Volume control)
17.2-5 |  | SIMV
17.2-6 |  | pressure support, advanced modes
17.3 | Monitor parameter readings
17.3-1 |  | Volume
17.3-2 |  | Pressure (Peak , Peep)
17.3-3 |  | Oxygen
17.3-4 |  | Frequency
17.3-5 |  | I : E ratio
17.4 | Monitor graphic waveforms | Volume
Pressure
17.5 | Technical data :
17.5-1 | Respiration frequency | 4-60 bpm.
17.5-2 | Inspiration flow | 0-180 L/min
17.5-3 | Tidal volume | 5-1500 ml
17.5-4 | I : E ratio | Please Specify
17.5-5 | pressure limit, cm H2O | Adjustable
17.5-6 | PEEP | 0-20 cm H2O or better
18 | Audible and Visual Alarms
18.1 |  | Low/ high pressure
18.2 |  | Low O2 concentration
18.3 |  | Low/ high Volume
18.4 |  | O2 inlet supply failure
18.5 |  | Alarm interruption 60 sec. with rest function
18.6 |  | battery alarm
18.7 |  | Power failure alarm
19 | Certification from the Manufacturer:
19.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
19.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
19.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
19.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
19.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
19.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
19.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
19.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
20 | Life Span for Machine: | 10 Years
21 | Maintenance:
21.1 |  | 2 years free maintenance.
21.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
21.3 |  | Application Software and Interface Connection Included.
21.4 |  | Spare Parts list with code NO
22 | Accessories:
22.1 |  | Anesthesia masks Resuable (all sizes)
22.2 |  | Sphygnomanometer
22.3 |  | Re-breathing bags Resuable (all sizes)
22.4 |  | Magil circuit
22.5 |  | Jackson-read modification of Ayri’s T piece
22.6 |  | O2 sensore
22.7 |  | All needed accessories to insure full use for adults, pediatrics and neonat.
23 | Self diagnostic and error / calibration message | Included
24 | Power Supply: | 100 - 240 V, 50 / 60Hz
Internal rechargeable battery
25 | Training: | User /Nurses training, by Specialist from the Supplier.
26 | Capnography/CO2 Monitoring: | Including anesthesia machine or patient monitor of Anesthesia machine
27 | Other Specification: | Please Specify.
OT-1 - 1 | Item Description | Patient Monitor of Anesthesia Machine
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the company.
6 | Display size and Type: | 12-15 Inches Or Better. Touch Screen
High resolution multi-color display
Medical type (preferable) or supported with isolation transformer
7 | Displayed information: | Min. 6 vital waveforms, cab be colored separately
Numeric data for the measured vital parameters
Vital parameters (72) hrs trends
Vital parameters alarms audio visual
8 | Defibrillation Protection: | Available for the measured vital parameters
9 | Required vital parameters:
9.1 | ECG:
9.1-1 | Leads: | 12 leads facility
9.1-2 | Values: | ECG lead waveform, label, HR gain as min.
9.1-3 | HR Range: | 30-200 bpm
9.1-4 | Alarms: | Leads off, Hi + Low HR
9.1-5 | Gain: | 5, 10, 20 mm/mV
9.1-6 | Preferable items: | ST, Arrhythmia, cascade ECG
9.2 | Respiration:
9.2-1 | Technology: | Impedance
9.2-2 | Values: | Respiration waveform, R.R.
9.2-3 | R.R. Range: | 5-80 bpm
9.2-4 | Alarms: | Hi & low R.R.
9.2-5 | Apnea Alarm: | 15-25 sec. Preferable
9.3 | NIBP:
9.3-1 | Technology: | Oscillometric
9.3-2 | Values: | SYS , DIA , MEAN , P.R.
9.3-3 | Modes: | Auto, Manual
9.3-4 | Cuff Pressure Range:
9.3-4-1 |  | Up to 250mmHg Adult
9.3-4-2 |  | Up to 200mmHg Pediatric
9.3-4-3 |  | Up to 150mmHg Neonate
9.3-4-4 | Alarms: | Hi and low, SYS, DIA
9.4 | SPO2:
9.4-1 | Values: | SPO2 waveform, SPO2% , P.R.
9.4-2 | SPO2 range | 50 - 100 %
9.4-3 | Pulse Rate Range: | 25-200ppm
9.4-4 | Alarms: | Hi & low SPO2 + P.R. sensor off
9.5 | Temperature:
9.5-1 | Values: | T1 and or T2
9.5-2 | Range: | 30-45°C
9.5-3 | Alarm: | Hi & low
9.6 | ETCO2:
9.6-1 |  | Should be Main Stream capnography with display of CO2 and digital Values of EtCO2, FiCO2 & RR.
9.6-2 | EtCO2 Range : | 0-99 mmHg.
9.6-3 | FiCO2 : | 0 to 20 mmHg.
9.6-4 | Flow rate : | 50ml/min.
9.7 | Invasive BP:
9.7-1 |  | Should have 2 channel Invasive Blood pressure (IBP) measurement.
9.7-2 |  | Should have waveform IBP1 and IBP2.
10 | Certification from the Manufacturer:
10.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
10.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
10.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
10.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
10.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
10.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
10.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
10.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
11 | Life Span for Machine: | 10 Years
12 | Maintenance:
12.1 |  | 2 years free maintenance.
12.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
12.3 |  | Application Software and Interface Connection Included.
12.4 |  | Spare Parts list with code NO
13 | Required Accessories:
13.1 |  | 1 X 3/5 Leads wire ECG cable Adult.
13.2 |  | 1 X 3/5 Leads wire ECG cable Pediatric.
13.3 |  | 1 X NIBP reusable cuff large Adult
13.4 |  | 1 X NIBP reusable cuff pediatric
13.5 |  | 1 X NIBP reusable cuff Infant
13.6 |  | 1 X NIBP connection hoses
13.7 |  | 1 X SPO2 reusable finger probe + Extension cable adult.
13.8 |  | 1 X Temperature reusable sensor skin type for , pediatric and Infant
13.9 |  | 1 X EtCO2 module with all accessories. ( Cable-Sensor)
13.1 |  | 2 X IBP Reusable Transducers with cable
13.11 |  | 10 X IBP Disposable pressure transducers
13.12 |  | Operation Manual
13.13 |  | Service Manual
13.14 |  | Spare parts list with code NO.
14 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
14.1 | Internal back-up battery | Operating time, hr : ≥2
15 | Other Specification: | Please specify.
OT-2 | Item Description | Operation Light (Ceiling) LED
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the company.
6 | Light Source, Type: | LED
7 | Number of Lightheads: | 2
8 | Number of Bulbs or Elements: | Specify
9 | Volts: | 12 V or 24 V
10 | Life, hr: | ≥50,000 Hr
11 | Color Temperature, K: | 3,000-5,000
12 | Color Adjustable: | Required
13 | Color Rendering index: | >90
14 | R9 Value: | Required
15 | Field size, cm Diameter: | ≥20
16 | Focal length, cm: | ≥ 70 cm
17 | Illumination level, maximum lux at 1 m: | 140.000 -160.000 Lux at 1.0m
18 | Mounting: | Ceiling
19 | Controls : | Required
19.1 | Dimmer: | With Dimmer
19.2 | Focus: | Adjustable
19.3 | Field size: | Adjustable
19.4 | Alarms: | main lides failure indicator
20 | On Sterile Handle: | On sterile handle sterilizable.
21 | Rotation: | 360
22 | Vertical adjustment range, cm: | ≥80
23 | Maximum irradiance at 1 m, W/m2: | <1,000
24 | Heat-to-light ratio, mW/m2.lux: | ≤3
25 | Lighthead Material: | high qualety Lighthead material
26 | Shadow Control: | Shadow less
27 | Battery or Emergency Backup: | Required (2-4) Hours Or Online UPS 3KVA
28 | Sterilizable Handle: | Required
29 | Certification from the Manufacturer:
29.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
29.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
29.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
29.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
29.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
29.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
29.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
29.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
30 | Life Span for Machine: | 7 Years
31 | Maintenance:
31.1 |  | 2 years free maintenance.
31.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
31.3 |  | Application Software and Interface Connection Included.
31.4 |  | Spare Parts list with code NO
32 | Required Accessories:
32.1 |  | 4 x Sterilizable handle
32.2 |  | 1 x Operation Manual
32.3 |  | 1 x Service Manual
33 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
34 | Other Specification: | Please specify.
OT-3 | Item Description | Operating Table "Major''
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the company.
6 | Operating Table type: | Electric & Hydrolic
7 | Maximum weight capacity, kg: | Static: 200-250
Articulated: ≥220
8 | Table Positions: | Specify
9 | Degrees From Horizontal: | Specify
| Trendelenburg: | ≥35
10 | Reverse Trendelenburg: | ≥30
11 | Lateral tilt: | ≥25
| Vertical range, cm: | 60-100 height-adjustable
12 | Table Section: | ≥5
13 | Number of Core Sections: | ≥3
| Degrees from Horizontal: | Specify
14 | Back Section: | 70 to30
15 | Foot/leg Section: | 20 to 90
| Head Section: | 45to85
16 | Controls: | Hand or foot and Remote Control
17 | Siderail: | Including.
| C-arm accessible: | Including.
18 | Base attachment: | Including.
19 | Column Housing: | Including.
| Conductive Casters: | Including.
20 | Caster Lock: | Including.
21 | Certification from the Manufacturer:
21.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
21.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
21.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
21.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
21.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
21.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
21.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
21.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
22 | Life Span for Machine: | 7 Years
22.1 | Maintenance:
22.2 |  | 2 years free maintenance.
22.3 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
22.4 |  | Application Software and Interface Connection Included.
22.5 |  | Spare Parts list with code NO
23 | Accessories:
23.1 |  | Anesthesia screen L type with clamp 1 piece
23.2 |  | Leg Support with clamp 2 pieces
23.3 |  | Arm Support with clamp 2 pieces
23.4 |  | Body Support 1 piece
23.5 |  | IV rod, 1 set, adjustable height with clamp
23.6 |  | Shoulder rest / support with pad, 2 pieces
23.7 |  | Head Plate 1 piece
23.8 |  | Chest & waist support
24 | Maintenance: | 2 years free maintenance.
25 | Other Specification: | Please specify.
OT-4 | Item Description | Electro Surgical Unit 400 W
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Generator Type: | Please Specify.
7 | Frequency: | High Frequency.
8 | Modes:
8.1 | Monopolar Mode: | Cut - Coagulate
8.1-1 | Cut Mode: | Low cut: maximum power of 400w, adjust output power Pure cut: maximum power of 300w, adjust output power Blend mode: maximum power of 300w, adjust output power
8.1-2 | Coag Mode: | Maximum 200-300 watts Please specify The Modes.
8.2 | Bipolar mode: | Please specify The Modes.
8.3 | Self-test Mode: | Including.
8.4 | Power Setting Displayed: | Including.
8.5 | Cooling : Convection | Including.
9 | Alarms: | Visual - audible
9.1 | For Plate Continuity: | Including.
9.2 | Excess Power: | Including.
9.3 | Internal Error Alarm: | Including.
9.4 | Plate voltage alarm: | Including.
9.5 | Patient Electrode Safety System: | Including.
9.6 | Defeatable-Inaudible Volume: | Including.
10 | Certification from the Manufacturer:
10.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
10.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
10.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
10.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
10.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
10.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
10.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
10.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
11 | Life Span for Machine: | 10 Years
12 | Maintenance:
12.1 |  | 2 years free maintenance.
12.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
12.3 |  | Application Software and Interface Connection Included.
12.4 |  | Spare Parts list with code No
13 | Accessories | Supplied with :
13.1 | Monopolar Footswitch: | Supply 1 pice
13.2 | Bipolar Footswitch: | Supply 1 pice
13.3 | Reusable hand switching Pencil: | Supply 2 pice
13.4 | Reusable Patient Plate: | Supply 2 pice
13.5 | Bipolar Forceps: | Supply 2 pice
13.6 | Forceps Cord: | Supply 2 pice
13.7 | Universal Adaptor: | Supply 1 pice
13.8 | Trolley: | Included Original Type
14 | Training: | User /Nurses training, by Specialist from the Supplier.
15 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
16 | Other Specification: | Please specify.
OT-5 | Item Description | Suction Machine (Mobile)
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Flow rate L/ Min: | Not less than approx. 30 liter / min.
7 | Vacuum , mmHg: | Not less than 700 mmHg.
8 | Vacuum regulate: | Manual kay adjust the vacuum volume to high or low.
9 | Bottles: | Two collecting jars (2-2.5) liters fitted with sterilizable rubber bung & fluid valve. Change-over block from jar to other
10 | Tube: | Suction tube between operation field and suction machine) reusable length 3m Material silicone, Autoclavable
11 | Castors: | The apparatus should be mounted on four anti-static tired rubber castors
12 | Filters: | The apparatus should be available with bacterial filters
13 | Certification from the Manufacturer:
13.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
13.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
13.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
13.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
13.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
13.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
13.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
13.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
14 | Life Span for Machine: | 7 Years
15 | Maintenance:
15.1 |  | 2 years free maintenance.
15.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
15.3 |  | Application Software and Interface Connection Included.
16 | Accessories: | 2 Spare bottles, 10 spare bacterial filters and 1 tube (Must be included in the price)
17 | Casters: | Four casters ,two with brakes
18 | Warranty: | 2 years
19 | Power Supply: | 100 to 240 V ~ ±10%, 50 Hz
20 | Other Specification: | Please specify.
OT-6 | Item Description | Anesthesia Cart
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
5.2 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Cart Dimension:
6.1 | Height: | >100 cm
6.2 | Depth: | 60 - 70 cm
6.3 | Width: | 50 -60 cm
7 | Drawers: | At Least 5 Drawers Different Sizes.
8 | Drawer Dimension:
8.1 | Height: | 6 - 8 cm
8.2 | Depth: | 40 -50 cm
8.3 | Width: | 50 -60 cm
9 | Castors: | Four casters ,two with brakes
10 | Accessories:
10.1 |  | Height adjustment I.V. Pole
10.2 |  | Divided Drawer Tray
10.3 |  | Tilt Bin Organizer
10.4 |  | Sharps Container
10.5 |  | Collapsible Working Table.
10.6 |  | Trash Cans
11 | Life Span for Machine: | 7 Years
12 | Warranty: | Minimum of 2 years
13 | Other Specification: | Please specify other specification
OT-7 | Item Description | Emergency Cart. (Crash Cart with Complete Access. )
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
5.2 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | General Description: | Emergency Cart, with accessories.
7 | Technical specifications: | Emergency trolley or (Emergency Crash Cart) is a set of trays/drawers/shelves on wheels used in hospitals for transportation and dispensing of emergency medication/equipment at site of medical/surgical emergency for life support protocols potentially to save a patient’s life.
8 | Constructions: | Stainless steel top and shelf & equipped with 4 - 6 removable coloured bins made of moulded plastic.
9 | System Configuration: | Emergency trolley (Emergency Crash Cart) with removable coloured bins, storage units, fitted with oxygen cylinder holder and electric lamp holder and four swivels castors.
10 | Dimensions: approx. | 960mm L x 500mm W x 1545mm H.
11 |  | Lockable storage units – 3 drawers (stainless steel or moulded plastic). Wood or wood laminate construction drawers are NOT acceptable.
12 |  | To be fitted with stainless steel, height adjustable, twin hook/loop, IV pole assembly
13 | Castors: | Fully, 360 deg. swivel castors/wheels, size 125mm dia with at least one castor/wheel to .have locking/brake mechanism.
14 | Shelves: | SS shelves 14 G thick stainless steel.
15 |  | Top shelf to have stainless steel guard rail approx.35mm above surface.
16 |  | Fitted with epoxy powder coated oxygen cylinder holder and electric lamp holder with clamp and cardiac massage board.
17 |  | Must be capable of carrying ECG Monitor/defibrillator and a suction apparatus.
18 | Accessories, spares and consumables: | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
19 | Life Span for Machine: | 7 Years
20 | Warranty: | Minimum of 2 years
21 | Other Specification: | Please specify Other Specification
OT-8 | Item Description | General Surgical Set Large
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Designed: | Heavy Duty.
5.1 | Quality: | High Quality.
5.2 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
5 | Each Set Contents of :
5.1 |  | 2 X Bowl Lotion 250 mm
5.2 |  | 2 X Galipot/Cup S/S 180 ml
5.3 |  | 2 X Galipot/Cup S/S 250 ml
5.4 |  | 2 X Tray Kidney shape with Lid S/S 200 X 30 mm
5.5 |  | 2 X Tray Kidney shape with Lid S/S 300 X 50 mm
5.6 |  | 20 X Forceps artery Kocher COF BJ.S/S 180mm.
5.7 |  | 15 X Forceps Artery Spencer Wells CO Flat (bj).S/S 200 mm
5.8 |  | 6 X Forceps Dissecting Plain serrated jaws S/S 150 mm
5.9 |  | 1 X Forceps Dissecting Plain serrated jaws S/S 250 mm
5.1 |  | 1 X Forceps Dissecting narrow, serrated jaws S/S 200 mm
5.11 |  | 3 X Forceps Dissecting Lane's 2x3 teeth 180 mm
5.12 |  | 1 X Forceps Dissecting Waugh, fine serr.jaws 1x2 Tth 200 mm
5.13 |  | 4 X Forceps Tissue Allis fine 4 x 5 Teeth BJ.S/S 150 mm
5.14 |  | 8 X Forceps Towel Holding cross Action S/S 100mm
5.15 |  | 4 X Forceps Towel Moynihan, Double Teeth (bj) 190mm
5.16 |  | 5 X Forceps Sponge Holding Rampley BJ S/S 240 mm
5.17 |  | 1 X Handle size 3 for Scalple blades size 10-15
5.18 |  | 1 X Handle size 4 for Scalple blades size 20-24
5.19 |  | 1 X Needle Holder Kilner's 165 mm
5.2 |  | 2 X Needle Holder Mayo Hegar BJ.S/S 180mm
5.21 |  | 1 X Scissors Blunt Point Straight SJ. S/S 180mm
5.22 |  | 1 X Scissors Mayo’s CO Flat 200mm S/S
5.23 |  | 3 X Scissors Mayo’s Straight 200mm S/S
5.24 |  | 1 X Scissors McIndoe's CO Flat 190mm S/S
5.25 |  | 1 X Retr.Deaver #1-Blade 25mm, Length 305mm
5.26 |  | 1 X Retr.Deaver #5-Blade 75mm, Length 305mm
5.27 |  | 2 X Retr. Lange beck, H/H 26x64 Blade, 210 mm Long
5.28 |  | 2 X Retr. Morris Double-ended (52x38)&(52x64) mm
5.29 |  | 1 X Retr. Walton, malleable copper 13mm length 305mm
5.3 |  | 1 X Needle Aneurysm S/S . Small 146 mm
5.31 |  | 4 X Pin Safety Mayo, forceps holding, 120mm
5.32 |  | 1 X Suction Tube McIndoe
5.33 |  | 1 XSuction Tube Poole's S/S 225mm
5.34 |  | 2 X Sterilizing and storing case St. St. with filters 260x150x50 mm
5.35 |  | 1 X Wire Cutting, Tungsten .Carbide inserts 180mm.
6 | Life Span for Machine: | 5 Years
7 | Warranty: | 2 years.
8 | Other Specification | Please specify.
OT-9 | Item Description | Ambubag for infant / Resuscitation Set (infant)
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
5.2 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Type: | Resuscitator, hand-operated, infant (body weight <5 Kg), set.
7 | Materials: | Latex-free
7.1 | Materials life: | All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions
7.2 |  | Made of transparent silicone rubber
7.3 | Capacity: | 280 ml
8 | Compressible Self-Refilling Ventilation Bag: | silicone rubber.
9 | Non-rebreathing Patient Valve with Pressure Limiting Valve: | polycarbonate/polysulfone/silicone
10 | Inlet Valve with Nipple for O2 Tubing: | polycarbonate/polysulfone
11 | Oxygen Reservoir Bag: | Plastic.
12 | Masks: | transparent silicone rubber.
13 | Accessories:
13.1 |  | Compressible self-refilling ventilation bag, capacity: 220-280ml
13.2 |  | Oxygen reservoir bag complete, capacity: 600ml
13.3 |  | mask 0 , face Round type, size infant 30-50 mm small
13.4 |  | mask, 1 face Round type, size infant 50-65 mm small
13.5 |  | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly
13.6 |  | Carrying case
14 | Life Span for Machine: | 7 Years
15 | Warranty: | Minimum of 2 years
16 | Other Specification: | Please specify Other Specification
OT-10 | Item Description | Phototherapy Unit
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Type: | Freestanding unit, Mobile on swivel castors , 2 (locks).
6.1 |  | Three or Four anti-static swivel castors, two at least with brakes
7 | Irradiance at 400 to 500nm, µw/cm2/nm: | Specify
8 | Angular Adjustments: | Required
9 | Bulbs Type: | LED
10 | Bulbs Number: | Specify
11 | Bulbs Expected Life: | >7000 hours
12 | Cooling Fan: | Required
13 | Hour Meter: | Required
14 | Timer: | Required
15 | Certification from the Manufacturer:
15.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
15.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
15.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
15.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
15.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
15.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
15.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
15.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
16 | Life Span for Machine: | 10 Years
17 | Maintenance:
17.1 |  | 2 years free maintenace.
17.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
17.3 |  | Application Software and Interface Connection Included.
18 | Accessories:
18.1 |  | 1 X Protective Cover.
18.2 |  | 2 X Neonatal Eye Mask.
18.3 |  | 3 X Fuses.
19 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
20 | Other Specification: | Please specify.
OT-11 | Item Description | Infant Incubator
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Type: | Mobile on swivel castors ,2 (locks).
7 | Temperature control range for air, °C: | 23-37°C with (0.1) °C increments and
7.1 |  | override up to 39 °C
8 | Temperature control range for skin, °C: | 32-37°C with (0.1) °C increments
9 | Temperature Display:
9.1 | Air , display type: | Please specify
9.2 | Range , °C: | 20-42 °C
10 | Temperature Display:
10.1 | Skin , Display type | LCD or similar not CRT with a size of min. 5" Or Better
10.2 | Range , °C: | 30-42 °C
11 | Alarms:
11.1 |  | High air temperature
11.2 |  | Low air temperature
11.3 |  | High skin temperature
11.4 |  | Low skin temperature
11.5 |  | Fan failure
11.6 |  | Sensor failure
11.7 |  | Power failure
12 | Front panel:
12.1 |  | Heater power indicator included
12.2 |  | Controls & switches should be visible, clearly labeled, protected against accidental changes and sealed against penetration of fluids
13 | Design:
13.1 |  | Compact design and heavy duty
13.2 |  | Backup over-heat protection included
13.3 |  | Supplemental humidity adjustable
13.4 |  | Air bacterial filter (changeable)
13.5 | Scale: | Yes
14 | Internal noise, dB: | ≤ 50 dB approx.
15 | Construction :
15.1 |  | Double wall
15.2 |  | Transparent hood for phototherapy and X-ray
15.3 |  | Movable tray with mattress
15.4 |  | Four antistatic swivel castors, two at least with breaks
15.5 |  | Drawer and instrument tray included
16 | Mattress:
16.1 |  | Adjustable to the trendelenburg and fowler position (Anti trendelenburg)
16.2 |  | Mattress height is adjustable up and down
17 | Ports:
17.1 |  | Hand ports
17.2 |  | Tubing ports:
17.3 |  | O2 inlet ports
18 | O2 Controllers: | Included ( servo Oxygen controllers)
19 | Backup Thermostat: | Required
20 | Heater Power Indicator: | Required
21 | X-Ray Tray: | Please specify
22 | Certification from the Manufacturer:
22.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
22.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
22.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
22.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
22.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
22.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
22.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
22.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
23 | Life Span for Machine: | 10 Years
24 | Maintenance:
24.1 |  | 2 years free maintenace.
24.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
24.3 |  | Application Software and Interface Connection Included.
25 | Accessories:
25.1 |  | 1 X Internal Scale Digital Type
25.2 |  | 1 X IV Pole
25.3 |  | 3 X Skin temperature
26 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
27 | Other Specifications: | Please specify.
OT-12 | Item Description | Syringe Pump W/holder
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Syring Pump: | Permits the use of standard disposable Syringes (sizes 10 – 50/60 ml) and extension sets, with auto syringe size sensing.
7 | Infusion Volume: | 0.1 to 999 ml in 0.1ml increments in lower ranges
8 | Infusion rate adjustable: | from 0.1 – 99.9 ml/hr in (0.1ml/hr steps)
9 | System and malfunction: | alarms with visual and audible alarm indicators
10 | Display:
10.1 |  | Digital display or better of volume infused. Infusion rate, alarm prompts etc.
10.2 |  | Purge/ bolus facility
10.3 |  | Volumetric accuracy ± 2% or less
10.4 |  | Occlusion Pressure limit (from 0 to 750 mmHg)
10.5 |  | Built-in self test and diagnostics
11 | Battery: | 4-6 Hrs.
12 | Model Release date: | Latest version
13 | Product Information:
13.1 |  | DPS (dynamic pressure sensor) Included
13.2 |  | Special feature of the product (other than mentioned on list)
13.3 |  | Accessories if any (to put machine/ system into operation)
13.4 |  | Consumable if any (to put machine/ system into operation)
14 | Training: | User /Nurses training, by Specialist from the Supplier.
15 | Certification from the Manufacturer:
15.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
15.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
15.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
15.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
15.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
15.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
15.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
15.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
16 | Life Span for Machine: | 10 Years
17 | Maintenance:
17.1 |  | 2 years free maintenace.
17.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
17.3 |  | Application Software and Interface Connection Included.
18 | Training: | User / Nurses training, by Specialist from the Supplier.
19 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
20 | Other Specification: | Please specify.
OT-13 | Item Description | Infusion Pump W/Holder
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Mode:
6.1 |  | Macro : 1-999 ml/hr in 1ml./hr increment
6.2 |  | Micro : 0.1-99.9 ml/hr in 0.1 ml/hr increment
7 | Volume to be infused: | 1.0 – 999.9 ml or No limit
8 | Infusion Set Used with the System: | should (open system) the user can use any infusion set
9 | Infusion Accuracy:
9.1 |  | Presettable rate
9.2 |  | Not Less Than ±2%
10 | Secondary Function:
10.1 |  | Free flow protection
10.2 |  | Front panel key pad lock
10.3 |  | Built-in self test and diagnostic
10.4 |  | System and malfunction alarms with audible and visual alarm indicators
10.5 |  | Silence or reset of audio alarms
10.6 |  | Air line detect system
11 | Message Display for Infusion Status: | Alarms, error messages etc.
12 | Battery Pickup: | 4-6 Hrs.
13 | Automatic Flow:
13.1 |  | Rate calculation
13.2 |  | Drug calculation & drug dose programs
13.3 |  | Multi-dose capability
13.4 |  | Rate-volume and volume timer programming
14 |  | Supplied With Spare Infusion Set For Adult & Pediatric. (each) 500 Set
15 | Certification from the Manufacturer:
15.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
15.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
15.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
15.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
15.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
15.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
15.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
15.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
16 | Life Span for Machine: | 10 Years
17 | Maintenance:
17.1 |  | 2 years free maintenace.
17.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
17.3 |  | Application Software and Interface Connection Included.
18 | Training: | User / Nurses training, by Specialist from the Supplier.
19 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz .
20 | Other Specification: | Please specify.
OT-14 | Item Description | Pulse Oximeter
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Display: | LCD Or LED. To Show , SpO2, pulse rate, pulse strength and/or signal, low battery.
7 | Type: | Desktop (Portable Type)
8 | SpO2 Range, %: | 70-100.
8.1 | Accuracy, %: | ±3.
9 | Pulse Rate bpm: | 30-240.
9.1 | Accuracy: | ±5%.
10 | Perfusion Index and/or Signal Strength Indicator. | Required, continuous monitors; preferred, spot check monitors.
11 | Response Time Sec: | Depends on application.
12 | Alarms :
12.1 | Audible and visual: | High/low SpO2 and pulse rate, sensor off, sensor failure, low battery
12.2 | Visual only: | High/low SpO2 and pulse rate, sensor off, low battery
13 | Alarm Override: | Preferred
13.1 | Reactivation method: | Automatic or manual
13.2 | Volume Control: | Required
14 | Self -Test Mode: | Required
15 | Sensor Types:
15.1 |  | Reusable Type
15.2 | Patient range: | Adult, pediatric, infant, neonate
15.3 | Cable length, m (ft): | 1.5 or more
16 | Battery:
16.1 | Type: | Rechargeable lithium-ion / or better.
16.2 | Life, hr: | 6-12 Hr
16.3 | Recharge time, hr: | ≤3.
17 | Certification from the Manufacturer:
17.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
17.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
17.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
17.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
17.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
17.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
17.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
17.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
18 | Life Span for Machine: | 10 Years
19 | Maintenance:
19.1 |  | 2 years free maintenace.
19.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
19.3 |  | Application Software and Interface Connection Included.
20 | Required Accessories:
20.1 |  | 1 X SpO2 Adult Resuable Sensor
20.2 |  | 2 X SpO2 Neonatal Resuable Sensor
20.3 |  | 1 X SpO2 Pediatric Resuable Sensor
20.4 |  | 1 X Charger
20.5 |  | 1 X Carying Bag
21 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz.
22 | Battery: | Rechargeable battery with built in charger with indicators for charging and low battery
23 | Other Specification: | Please specify.
ICU-1 | Item Description | Ventilator Machine
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the company.
6 | Ventilator System:
6.1 |  | Microprocessor controlled used for Pediatric,Adult and Neonate.
6.2 |  | Active Humidifier.
6.3 |  | Compressor or turbine >60000 hour - full 7 years warranty
6.4 |  | Nebulizer
7 | Ventilation mode | volume and pressure type of breaths , Patient types Adult & pediatric.and Neonate
7.1 |  | BIPASIC , VG and Bi level ventilation modes ,PC-CMV, PC-BIPAP1 / SIMV, PC-SIMV, PC-AC, PC-APRV, PC-PSV,VC-MMV, A/C, CPAP/PEEP, VC-AC,VC-CMV, VC-SIMV, SPN-CPAP/PS, SPN-CPAP/VS, SPN-CPAP, SPN-PPS)and newer modes for Adult & pediatric.and Neonate
7.2 | Invasive Ventilation Modes: | Please specify.
7.3 | Non-invasive ventilation Modes | Required With all Modes.
8 | Controls /Setting Ranges:
8.1 | Tidal volume, mL | 2-2000 ml Or Better
8.2 | Respiratory rate, breaths/min | 1-150
8.3 | Trigger mechanism | Pressure, flow
8.4 | FiO2, % | 21-100
8.5 | Inspiratory flow rate, L/min | 0-180
8.6 | Flow pattern/waveform adjustment | Required
8.7 | Inspiratory pressure, cm H2O | 0-80
8.8 | IE ratio | Please specify.
8.9 | Sigh breath function | Preferred
8.1 | PEEP/CPAP, cm H2O | 0-20
8.11 | Pressure support, cm H2O | 0-40
8.12 | Leak compensation | Required
8.13 | Auto 100%/Increase O2 button | Required
8.14 | Control panel lock | Required
8.15 | Automatic ventilation mode | Required
9 | Special Functions:
9.1 | Pediatric application | Required
9.2 | Neontel application | Required
9.3 | Nebulizer | Required
9.4 | Apnea pack up ventilation | Required
9.5 | Pause (plateau ) | Required
10 | Graphic waveform , Trends & Loops:
10.1 |  | Flow
10.2 |  | Volume
10.3 |  | Pressure
11 | Ventilation Display / Monitored: | Multicolor LCD / LED display
11.1 |  | Peak inspiratory pressure
11.2 |  | Mean airway pressure
11.3 |  | PEEP pressure
11.4 |  | Tidal volume
11.5 |  | Minute volume
11.6 |  | Spontaneous minute volume
11.7 |  | FiO2 (analyzed %)
11.8 |  | Respiratory rate
11.9 |  | Inspiratory time
11.1 |  | Expiratory time
11.11 |  | IE ratio
12 | Alarm limits (adjustable):
12.1 |  | Low/high pressure limit
12.2 |  | Low O2
12.3 |  | Low / high PEEP / CPAP pressure
12.4 |  | Low / high tidal volume
12.5 |  | Low / high minute volume
12.6 |  | Low / high respiratory rate
12.7 |  | O2 Concentration : 18-100 %
12.8 |  | Apnea
12.9 |  | All alarms audible and visible
13 | Equipment Alarms:
13.1 |  | Power failure
13.2 |  | Low battery
13.3 |  | Disconnection
13.4 |  | Gas supply failure
13.5 |  | Flow sensor alarms
13.6 |  | Vent inoperative
13.7 |  | Self-diagnostics
13.8 |  | All alarms: audible and visible
14 | Patient Alarms:
14.1 |  | Low/high FiO2
14.2 |  | Low minute volume
14.3 |  | High minute volume
14.4 |  | Low inspiratory pressure
14.5 |  | High pressure
14.6 |  | Loss of PEEP
14.7 |  | Apnea
14.8 |  | Continuous high pressure/occlusion
14.9 |  | Inverse IE ratio
14.1 |  | High respiratory rate
14.11 |  | High PEEP
14.12 |  | Breathing circuit disconnect
15 | Reporting (vent alarms and patient status):
15.1 | View reports on display | Required
15.2 | Save data to USB | Required
15.3 | Send data via network | Required
16 | Inlet gas supply pressure (Air/O2): | (30-80) PSI
17 | On-Board Air Compresspr Or Turbine | Please specify
17.1 | Compressor for ventilator: | Medical type, high quality and heavy duty
17.2 |  | Original and compatible with the ventilator
17.3 |  | Water and oil free
17.4 |  | Low noise ( > 60 dB)
17.5 |  | Automatic compressor backup system .(Air compressor will switch on automatically when inlet pressure is failed )
17.6 |  | Easy accessible
17.7 | Turbine: | High Quality Turbine >60000 hour - full 7 years Warranty Or more.
18 | Display : | Led, touchscreen
19 | Display size, (in) : | min 15 in
20 | Certification from the Manufacturer:
20.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
20.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
20.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
20.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
20.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
20.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
20.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
20.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
20.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
21 | Life Span for Machine: | 10 Years
22 | Maintenance:
22.1 |  | 2 years free Maintenance.
22.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
22.3 |  | Spare Parts list with code NO
23 | Accessories:
23.1 |  | 1x O2 sensor (Permanent Type)
23.2 |  | 1x Active Humidifier
23.3 |  | 2 x Reusable patient circuit for pediatric
23.4 |  | 2 x Reusable patient circuit for adult
23.5 |  | 2 x Reusable patient circuit for Neonate
23.6 |  | 1 x Resuable Test lung for Adult , Pediatric and Neonate
23.7 |  | 1 x Resuable Test lung for Pediatric and Neonate
23.8 |  | 3 x Disposable patient circuit for pediatric
23.9 |  | 3 x Disposable patient circuit for adult
23.1 |  | 3 x Disposable patient circuit for Neonate
23.11 |  | Nebulizer Build-in
23.12 |  | All needed accessories for full pediatric & adult and Neonatal use
24 | Training | User /Nurses Training, by Specialist from the Supplier. And Biomedical Engineers
25 | Internal back-up battery | Operating time, hr : ≥2
26 | Self diagnostic error / calibration massage | Included
27 | Power supply | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
28 | Other specification | Please specify.
ICU-2 | Item Description | Patient Monitor With Holder
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the company.
6 | Display size and Type: | 12-15 Inches Or Better. Touch Screen
High resolution multi-color display
Medical type (preferable) or supported with isolation transformer
7 | Displayed information: | Min. 6 vital waveforms, cab be colored separately
Numeric data for the measured vital parameters
Vital parameters (72) hrs trends
Vital parameters alarms audio visual
8 | Defibrillation Protection: | Available for the measured vital parameters
9 | Required vital parameters:
9.1 | ECG:
9.1-1 | Leads: | 12 leads facility
9.1-2 | Values: | ECG lead waveform, label, HR gain as min.
9.1-3 | HR Range: | 30-200 bpm
9.1-4 | Alarms: | Leads off, Hi + Low HR
9.1-5 | Gain: | 5, 10, 20 mm/mV
9.1-6 | Preferable items: | ST, Arrhythmia, cascade ECG
9.2 | Respiration:
9.2-1 | Technology: | Impedance
9.2-2 | Values: | Respiration waveform, R.R.
9.2-3 | R.R. Range: | 5-80 bpm
9.2-4 | Alarms: | Hi & low R.R.
9.2-5 | Apnea Alarm: | 15-25 sec. Preferable
9.3 | NIBP:
9.3-1 | Technology: | Oscillometric
9.3-2 | Values: | SYS , DIA , MEAN , P.R.
9.3-3 | Modes: | Auto, Manual
9.3-4 | Cuff Pressure Range:
9.3-4-1 |  | Up to 250mmHg Adult
9.3-4-2 |  | Up to 200mmHg Pediatric
9.3-4-3 |  | Up to 150mmHg Neonate
9.3-4-4 | Alarms: | Hi and low, SYS, DIA
9.4 | SPO2:
9.4-1 | Values: | SPO2 waveform, SPO2% , P.R.
9.4-2 | SPO2 range | 50 - 100 %
9.4-3 | Pulse Rate Range: | 25-200ppm
9.4-4 | Alarms: | Hi & low SPO2 + P.R. sensor off
9.5 | Temperature:
9.5-1 | Values: | T1 and or T2
9.5-2 | Range: | 30-45°C
9.5-3 | Alarm: | Hi & low
9.6 | ETCO2:
9.6-1 |  | Should be Main Stream capnography with display of CO2 and digital Values of EtCO2, FiCO2 & RR.
9.6-2 | EtCO2 Range : | 0-99 mmHg.
9.6-3 | FiCO2 : | 0 to 20 mmHg.
9.6-4 | Flow rate : | 50ml/min.
10 | Certification from the Manufacturer:
10.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
10.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
10.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
10.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
10.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
10.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
10.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
10.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
10.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
11 | Life Span for Machine: | 10 Years
12 | Maintenance:
12.1 |  | 2 years free Maintenanc.
12.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
12.3 |  | Application Software and Interface Connection Included.
12.4 |  | Spare Parts list with code NO
13 | Required Accessories:
13.1 |  | 1 X 3/5 Leads wire ECG cable Adult.
13.2 |  | 1 X 3/5 Leads wire ECG cable Pediatric.
13.3 |  | 1 X NIBP reusable cuff large Adult
13.4 |  | 1 X NIBP reusable cuff pediatric
13.5 |  | 1 X NIBP reusable cuff Infant
13.6 |  | 1 X NIBP connection hoses
13.7 |  | 1 X SPO2 reusable finger probe + Extension cable adult.
13.8 |  | 1 X Temperature reusable sensor skin type for , pediatric and Infant
13.9 |  | 1 X EtCO2 module with all accessories. ( Cable-Sensor)
13.1 |  | 1 X Original Holder In The Wall
13.11 |  | Operation Manual
13.12 |  | Service Manual
13.13 |  | Spare parts list with code NO.
14 | Training: | Users /Nurses and Biomedical Engineers Training, by Specialist from the Supplier.
15 | Power Supply: | 100 to 240 V ±10%, 50 Hz.
15.1 | Internal back-up battery | Operating time, hr : ≥2
16 | Other Specification: | Please specify.
ICU-3 | Item Description | ECG machine 12 ch. With Trolley
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required Certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty Designed, New Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | High Quality, Safety, Ease of Use
5.4 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | System: | Fully automatic, microprocessor controlled, channel resting electrocardiograph, mains and battery operated, built-in LCD
7 | Modes: | Suitable for Adult/Pediatric/Neonate modes, and easy to all functions and parameter settings. Alarm mute On/Off recorder, Built-in alarm mute, on/off freeze, on/off recorder, Leads Switching, automatic signal Calibration, and Leads-off indicator, Automatic Switch off: After 5 min inactive
8 | Sensitivity: | 2.5, 5, 10, 20 mm/mV.
9 | Frequency: | Range for diagnostic: 0.67-150 Hz, range for filtering: Any line frequency, low and muscle artifact/high frequency.
10 | Recording: | Channels: Multi channels, ECG leads: 12 standard leads, method: Automatic/manual, speed: 35, 65 mm/sec, automatic printout on standard paper printer if printer is connected, printout on internal thermal printer for realtime manual mode.
11 | Measurements: | PR, PQ, QT, ATC, P, QRS, T, HR; others by requirements and clinician preference. PR, PQ, QT, ATC, P, QRS, T, HR; others by requirements and clinician preference, Arrhythmia analysis.
12 | Visual Indicators: | Visual Indicators: Loose contact, lead fail, low battery, system status, artifacts and anthers.
13 | Patient leakage current: | Less than 10µ A.
14 | Battery | Included. At Least 4 Hrs.
15 | Certification from the Manufacturer:
15.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
15.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
15.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
15.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
15.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
15.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
15.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
15.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
15.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
16 | Life Span for Machine: | 10 Years
17 | Maintenance:
17.1 |  | 2 years free Maintenanc
17.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
17.3 |  | Application Software and Interface Connection Included.
18 | Accessories: | 3 x Roll Paper
18.1 |  | 1 X High Quality Trolley With Arm.
18.2 |  | 1 X Case, carrying.
18.3 |  | 2 X Packs, thermosensitive. 50mmx20m.
18.4 |  | 2 X Electrodes, Limb. clamp type.
18.5 |  | 2 Kits X Electrodes, Suction.
18.6 |  | 200 Pcs X Electrodes, Resting. Disposable.
18.7 |  | 10 X Adaptors, Clip.Alligator.4mm.
18.8 |  | 2 X 250g Gel, electroconductive.
18.9 |  | 2 X Patient cable. Banana plug.
19 | Power supply | 100 to 240 V ~ ±10%, 50/60 Hz.
20 | Other Specification: | Please specify.
ICU-4 | Item Description | Electric I.C.U. Bed
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of Origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | CE Mark: | Required Certificate.
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, Europe Union, Japan, and USA
5 | Designed: | Heavy duty Designed, New Model.
5.1 | Quality: | High Quality.
5.2 | Supplier: | the Supplier is required to be a Certified Distributor or Agent of the Company.
6 | Type: | Electric Critical Care Beds
7 | Patient Controls:
7.1 | Location: | Clearly visible with symbols
7.2 | Functions: | 5
8 | Nurse Controls: | Positioning, nurse call
8.1 | Patient control lockouts: | Required
8.2 | Full-low Indicator: | Required
8.3 | CPR control | Required
8.4 | Trendelenburg head section Indicator: | Required
9 | Automatic Contour :
10 | Retractable : | Required
11 | Maximum Patient Wehght, kg: | ≥ 200 kg
12 | Bariatric Features:
13 | Radiographic/Fluoroscopic Features: | Diolucent window for fluoroscopy/radiology beds
14 | Siderail Length:
14.1 | Length, cm (in): | Please Specify
14.2 | Percentage of overall: | Please Specify
15 | Casters:
15.1 | Diameter, cm: | 12.5
15.2 | Functions: | Swivel, conductive, ≥2 locking
16 | Central Brake System : | Required
17 | Bumpers: | Required
18 | Removable Headboard: | Required
19 | Removable Fottboard: | Required
20 | IV Pole: | ≥1 on each side
20.1 | Mounts: | Required
21 | Overall Dimensions:
21.1 | L x W, cm: | Approx. 210 x 100
21.2 | Height, cm: | Adjustable, ≥84
22 | Weight, kg: | 200 kg
23 | Battery Backup: | Yes
23.1 | Operating time, hr: | Please Specify
23.2 | Rechargeable: | Yes
24 | Power Supply: | 100 to 240 V ~ ±10%, 50/60 Hz Single phase
25 | Life Span for Machine: | 10 Years
26 | Warranty: | 2 Years
27 | Accessories:
27.1 |  | 1 X Mattres 10-15 cm (4 Sections)
27.2 |  | 1 X Over Bed Table
27.3 |  | 1 X Bedside Cabinet
28 | Other specification | Please specify.
ST-1 | Item Description | Bench Top Steam Sterilizer With Dryer (Autoclave)
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, European Union, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier | the supplier is required to be a certified distributor or agent of the company.
6 | Air Removal Method: | To facility requirements
7 | Chamber Material: | 314 Stainless Steel
7.1 | Type: | Class B
7.2 | Number of trays: | 2-3.
8 | Water Reservoir Capacity , L: | 24-40 L
9 | Preprogrammed Cycles, Type: | Wrapped, Unwrapped
10 | Cycle Number ,Type: | Please Specify
11 | Minimum Cycle Time ,Min: | Please Specify
12 | Automatic Cycle Shut off: | Included.
13 | Temperature Range , °C : | 100-134 °C
13.1 | Overheat shutoff: | Included.
14 | Pressure Purge: | Please Specify
15 | Controller Type: | Please Specify
15.1 | Programmable: | Included.
16 | Alarm Conditions:
16.1 | Indicator type: | Audible, visual
16.2 | Chamber pressure or temperature Failure: | Audible, visual
16.3 | Component Failure: | Audible, visual
16.4 | Cycle disruption or Failure: | Audible, visual
17 | Data Management:
17.1 | Recorder: | Included.
17.2 | Printer: | Included.
18 | Power Supply, VAC: | 110-230 50/60 Hz
18.1 | Consumption, Kw: | Please Specify
19 | Drayer Function: | Included. Vaccum Type
20 | Accessories:
20.1 |  | Steel container : plain basket with handles.
20.2 |  | Paper for printer Qty: 10 rolls
20.3 |  | Spare Part Heater 1 Pcs.
20.4 |  | Rubber Gasket for Door 1 Pcs.
21 | Certification from the Manufacturer:
21.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
21.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
21.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
21.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
21.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
21.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
21.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
21.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
21.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
22 | Life Span for Machine: | 10 Years
23 | Maintenance:
23.1 |  | 2 years free maintenance.
23.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
23.3 |  | Application Software and Interface Connection Included.
24 | Power Supply: | AC 110-240V, 50/60Hz single-phase electrical source.
24.1 |  | Compliant with IEC 60601.
24.2 |  | Built-in protections against over-voltage and over-current line conditions.
25 | Other Specification: | Please specify.
Company Name اسم الشركة |  |  | Signature & Stamp التوقيع و الختم
Representative Nameاسم ممثل الشركة
Designation المسمى الوظيفي
Date: التاريخ

## Technical Specification(LOT3) (hidden)
Annex B (LOT 3)- Technical Specification نموذج المواصفات الفنيه X RAY Digital 500 for Aslam HC - Hajjah
Technical Specifications
1 | Item Description | X RAY Digital 500 mA (HF)
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier: | the supplier is required to be a certified distributor or agent of the company.
6 | Type: | The Radiography System , should be of latest model and must have the latest technology, The system must belatest Generation, New model, Radiography System able for high load & hard work.
7 | General Description:
7.1 | X-Ray Generator:
7.2 |  | Generator should be of latest high frequency inverter technology for constant output and lowest radiation doses.
7.3 |  | Solid state high frequency (20 kHz or more) generator with minimum ripples having at least 3-5 kW output with ( PSU) Capacitor bank . Latest compact size generator assembly preferably integrated into table.
7.4 |  | Kv range: 40 -150 kV with 1 kV steps.
7.5 |  | Exposure time range: 1 millisecond (or less) to 5 seconds (or more).
7.6 |  | Digital Display of mA, kV, mAs on console panel
7.7 |  | Should have 500mA.
7.8 |  | More than 250 anatomical programmable radiography (APR) presets loaded for ease of use. Bidder to specify number of programs.
7.9 |  | Power input to be 230-240 VAC, 50 Hz with transformer provided by the supplier if the voltage is different, fitted with industrial plug.
7.1 |  | Resettable overcurrent protection shall be fitted with electromagnetic circuit breaker.
8 | X-Ray Tube and Collimator:
8.1 |  | Should be a high speed rotating anode dual focus tube of 2600 rpm or more compatible with the generator.
8.2 |  | Should have dual focal spots with the following focal spot size range: small focal spot size: 0.6 or better, large focal spot size: 1.2mm or better. Smaller focal size would be preferred.
8.3 |  | mA range: 10-500 milliampere or more.
8.4 |  | mAs range: 0.5-500 (or more)
8.5 |  | Tube anode heating capacity: at least 300KHU
8.6 |  | Should have a collimator with auto-off function.
8.7 |  | Incoming voltage indicator should be present.
8.8 |  | Manual shutter control collimator.
8.9 |  | Should have a multi leaf collimator having halogen/bright light source Or Led with auto shut provision for the light.
8.1 |  | Should have over load protection .
9 | X-Ray Table / Horizontal Bucky:
9.1 |  | Table top should be a carbon fiber top at least 220 cm (length) and 80 cm (width) Or Better
9.2 |  | Table top height (from ground) to be at least 65cm.
9.3 |  | Table top material to have low radiation absorption.
9.4 |  | The unit should be coupled to a horizontal table having floating table top with both longitudinal.
9.5 |  | It should have front pedals with electromagnetic locks for locking and releasing the table movements.
9.6 |  | The table should have a mobile bucky with a grid ratio of 12:1 (or better) at a focal distance of 115 cm. The bucky should be compatible with standard size cassette 35*43cm (14”x17”).
9.7 |  | Auto-centering of X-Ray tube over the bucky (in the transverse direction) after every exposure.
9.8 |  | It should have a weight bearing capacity of 200kg or more.
9.9 |  | Power input to be 220-240VAC, 50HZ
9.1 |  | Patient hand grips
10 | Vertical Tube Stand:
10.1 |  | Tube stand to be integrated with table, requires no wall/ceiling support.
10.2 |  | It should have manual locking for various movements .
10.3 |  | It should have movements in all directions.
10.4 |  | All movements should have electromagnetic brakes with fully counter balanced mechanism.
10.5 |  | It should have facility to display FFD/SID (Source to Image Distance) in vertical positions 150 cm or more, in horizontal position 180 cm or more.
10.6 |  | It should have provision of auto centering with the detector.
10.7 |  | Tube rotation at vertical axis and horizontal axis +/ - 180 degree.
10.8 |  | Cranio-caudal tube tilt (tilt along long axis of the table) to be -200 to +200 or better.
11 | Vertical Bucky Stand:
11.1 |  | The Bucky Should have Multi Movments, vertical and oblique positions with suitable movements allows for a complete range of exams from skull, skeletal body including spine, chest, bearing knee and ankle exams.
11.2 |  | Should have a vertical bucky with oscillating/moving grid for chest radiography (grid ratio - 10:1 or better).
11.3 |  | It should have provision to do chest radiography without grid.
11.4 |  | Could have automatic exposure control with at least 3 fields.
11.5 |  | The detector should be capable of rotating on its axis across +90 to -15 degrees.
11.6 |  | The vertical movement range should be 125cm or more with the lowest point (from cassette centre to ground) being not more than 55cm and the highest point being not less than 175 cm.
11.7 |  | The bucky should have electromagnetic lock that allow for easy positioning
12 | Certification from the Manufacturer:
12.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
12.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
12.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
12.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
12.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
12.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
12.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
12.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
12.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
13 | Life Span for Machine: | 10 Years
14 | Maintenance:
14.1 |  | 2 years free maintenance.
14.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
14.3 |  | Application Software and Interface Connection Included.
15 | System Configuration Accessories ,Spares and Consumables:
15.1 |  | 1 X Color LaserJet Printer for reporting
15.2 |  | 1 X Apron
16 | Training:
16.1 |  | one engineer and 2 users On how to use and perform periodic preventive maintenance of the device..
17 | Power Supply: | AC 110-240V, 50/60Hz single-phase electrical source.
17.1 |  | Compliant with IEC 60601.
17.2 |  | Built-in protections against over-voltage and over-current line conditions.
18 | Other Specification: | Please specify other specification
External DR System With Complete Accessories
No. | Technical Specifications | Minimum Requirement | Suppliers' Response | Remarks
1 | Manufacturer : | Please specify manufacturer and country of origin
2 | Model Number: | Please specify model number of the offered equipment
3 | Country Of origin: | Please specify Country of origin of the offered equipment
4 | Safety Standard:
4.1 | FDA Certificate Or CE Mark: | Required certificates
4.2 | If The Machine Have CE Certificate Product Needs To Have Market Clearance For Either: | Australia, Canada, EU, Japan, and USA
5 | Genaral Specification: | The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl.
5.1 | Designed: | Heavy duty designed, new Model.
5.2 |  | The machine shuld be of latest model and must have lattest technology
5.3 | Quality: | high quality, Safety, Ease of Use
5.4 | Supplier: | the supplier is required to be a certified distributor or agent of the company.
6 | Type: | The DR System , should be of latest model and must have the latest technology, The system must belatest Generation, New model, DR able for high load & hard work.
7 | 1. Digital Detector:
7.1 |  | One Detector ,We Can Installed in the radiological table and the vertical/chest bucky.
7.2 |  | The size of the detector should be 35 cm x 43 cm or more.
7.3 |  | The active matrix size should be 2800 X 2400 pixels or more at 140µm pitch.
7.4 |  | Should have a minimum image depth of 12-16 bit. Or Better
7.5 |  | Housing material: built in material resistant to blows and falls
7.6 |  | Interface: Ethernet . Cables and necessary input and output devices (USB) to connect it to the Computer
7.7 |  | The detector software shall be capable of operating with Windows 10 (OS) or higher (This shall be with licensed software for OS and Applications).
7.8 |  | Material : Preferred Made of amorphous selenium for high image quality. or Other high-Resolution Materials that are better than Required.
7.9 |  | With all accessories (extra batteries + battery charger) , Power Adapter , Control Box ,Bluetooth
8 | 2. DR Workstation (Image Acquisition,Image Processing):
8.1 |  | The digital workstation should be based on the latest high speed processors of at least 32 bit.
8.2 |  | It should have the capability of acquiring the image from the detector system.
8.3 |  | Should have preview time 5 seconds or better.
8.4 |  | The system should be ready DICOM interface and networking capability with RIS/HIS/PACS.
8.5 |  | Should provide for HL-7 compatible interface.
8.6 |  | Advance Post Processing Software with function: for sorting of patient image based on name, date, exam etc. using predefined parameters or user defined and stored image parameters.
8.7 |  | Correcting typographical in patient demographic module, in case RIS connection was down and manual data entry was done.
8.8 |  | Capability of changing R/L, Flipping, Rotating, Zooming, Collimating, annotating the incoming image.
8.9 |  | Workstation: one (1) latest Genaration system, Processor (Intel Core i7 or better): 3.4 GHz or better, with Windows 10.1 (OS) or higher, minimum 16 GB RAM, minimum 4 Tera-Byte Hard disk SSD, Medical grade 19'' Or 24 '' monitor supported by all necessary software for all the various DR functions. All the accessories like mouse, keyboard, power cable etc.
9 | 3. Dry Laser Camera/Dry -View Imaging Printer (film based) with the following:
9.1 |  | Print Images from DR workstation. In DICOM 3 format.
9.2 |  | Mechanism to print images to 8x10 and 10x12, 11x14, 14x17 film sizes (with minimum 2 universal tray online)
9.3 |  | Resolution > 500 DPI or more.
9.4 |  | Throughput: minimum 45-60 films per hour of size 14 x 17 in. (35 x 43 cm)
9.5 |  | Multiple Image and slide printing capability.
9.6 |  | Ethernet network compatible.
9.7 |  | Suitable online UPS with minimum 30-60 minutes backup time.
10 | Certification from the Manufacturer:
10.1 |  | That the bidder has the capability for corrective and preventive maintenance of the unit.
10.2 |  | That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded.
10.3 |  | Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer.
10.4 |  | Guaranteeing the availability of all spare parts for the next ten (10) years.
10.5 |  | That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model.
10.6 |  | That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract.
10.7 |  | Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment.
10.8 |  | Technical support from the manufacturer incase the agent or distributor doesn’t response when needed.
10.9 |  | The Supplier Company's Commitment to Calibration and Maintenance Required for the Device During the Maintenance Guarantee Period from The Beginning of Operating the device to the Delivery of the Device after the end of the free Warranty According to National and International Standards and Manfucturer Requirements.
11 | Life Span for Machine: | 10 Years
12 | Maintenance:
12.1 |  | 2 years free maintenace.
12.2 |  | Service Manual and Operation Manual {Hardcopy & Softcopy}
12.3 |  | Application Software and Interface Connection Included.
13 | Training:
13.1 |  | one engineer and 2 users On how to use and perform periodic preventive maintenance of the device..
14 | Power Supply: | AC 110-240V, 50/60Hz single-phase electrical source.
14.1 |  | Compliant with IEC 60601.
14.2 |  | Built-in protections against over-voltage and over-current line conditions.
15 | Other Specification: | Please specify other specification
Company Name اسم الشركة |  |  | Signature & Stamp التوقيع و الختم
Representative Nameاسم ممثل الشركة
Designation المسمى الوظيفي
Date: التاريخ